## 01_CONTENTS
Sheet | Purpose
01_CONTENTS | Index of primary results workbook sheets.
02_METADATA | Image-level metadata and rule-based image QC used to audit the analysis.
03_OVERVIEW | Final quantitative summary, selected method, reliability and external reference checks when provided.
04_RAW | Well-level analytical values used by the fitting pipeline.
05_REPLICATES_MEAN | Replicate-group medians, robust SDs and group QC flags.
06_FITTING | Calibration, standard-addition and unknown/CRM fit results.
07_METHOD_COMPARISON | Method ranking using only common score factors; expected values are external checks only.
08_LEGENDS | Definitions for all fields reported in this workbook and primary RGB figure.

## 02_METADATA
Field | Value | Notes
Analysis image source | original | original or in-memory resized image used for all measurements
Original image path | C:/Users/prova/OneDrive - unifi.it/Desktop/cromo/Piastra viola/Calibration Purple.jpg | source file selected by the user
Original image size | 1819 x 1204 px | width x height in pixels, when known
Analysis image size | 1819 x 1204 px | width x height in pixels actually analyzed
Resize scale | 1 | analysis_width / original_width, when a resized image is used
Initial image QC | WARNING | Rule-based result: FAIL if destructive=True; WARNING if quality_warning=True or borderline=True; otherwise OK. No image correction is applied.
Initial image QC class | usable_with_warnings | non_correctable if destructive=True; quality_warning if flatfield_span > 0.18 and no destructive defect; usable_with_warnings if only borderline defects; good if no flags.
Initial image QC messages | borderline blur | Messages list the threshold rules triggered by the measured fields below.
Image QC decision rules | thresholds | destructive = dead_channel OR saturation_fraction>0.003 OR saturation_all_channels_fraction>0.0008 OR specular_fraction>0.003 OR max_side<900 OR approx_roi_pixels<120; borderline = saturation_fraction>0.0005 OR specular_fraction>0.0005 OR blur_score<35; quality_warning = flatfield_span>0.18.
Flat-field span | 0.12461 | Flatfield_span = (P95 - P5) / median of the slow illumination field estimated from the grayscale image. Larger values indicate stronger illumination/background gradient.
Specular fraction | 0 | fraction of bright low-saturation pixels
Saturation fraction | 0 | fraction of pixels with at least one saturated channel
Dead channel | 0 | 1 if a channel is near-dead by mean/variance checks
Approx well pitch | 128.5202 | median distance between adjacent well centers, in pixels
Approx ROI pixels/well | 5039.5 | median number of pixels used in the final floor ROI mask per well
Blur score | 13.2961 | variance of Laplacian of the analysis image; lower values indicate stronger blur

## 03_OVERVIEW
Field | Value
Quantification | available
Reliability score | 80
Confidence class | HIGH
Reliability note | calibration and samples can be compared within the same plate
Empty-well QC status | ok
Selected method from rank | Signal_Green
Selected family | RGB
Ranking mode | calibration_only
Selected method score | 0.99525
R2 calibration | 0.99525
LOD | 0.21397
LOQ | 0.71325
Status notes | intraplate calibration present; empty-well QC: ok

## 04_RAW
Row | Col | Well | ID | Type | Conc | DF | MeanW_Red | MeanW_Green | MeanW_Blue | MeanBG_Red | MeanBG_Green | MeanBG_Blue | SignalT_Red | SignalT_Green | SignalT_Blue | PAbs_Red | PAbs_Green | PAbs_Blue | L | a | b | DeltaL | Deltaa | Deltab | DeltaE_ab | DeltaE_ab_chroma | CIELAB_ref_source | ImageWarning
A | 1 | A1 | 1 | C | 0 | 1 | 0.52344 | 0.58597 | 0.59894 | 0.5434 | 0.55819 | 0.64592 | 0.96327 | 1.0498 | 0.92727 | 0.016253 | -0.021099 | 0.032795 | 80 | -3 | -2 | 0 | 0 | 0 | 0 | 0 | zero_calibration | 0
A | 2 | A2 | 1 | C | 0 | 1 | 0.52952 | 0.59244 | 0.60548 | 0.54002 | 0.55925 | 0.64514 | 0.98056 | 1.0593 | 0.93853 | 0.0085245 | -0.025036 | 0.027552 | 80.3922 | -3 | -2 | 0.39216 | 0 | 0 | 0.39216 | 0 | zero_calibration | 0
A | 3 | A3 | 1 | C | 0 | 1 | 0.52344 | 0.58597 | 0.60548 | 0.53712 | 0.56002 | 0.6434 | 0.97454 | 1.0464 | 0.94107 | 0.011201 | -0.019677 | 0.026377 | 80 | -3 | -2 | 0 | 0 | 0 | 0 | 0 | zero_calibration | 0
A | 4 | A4 | 1 | C | 0 | 1 | 0.52344 | 0.57955 | 0.61207 | 0.5347 | 0.56048 | 0.6407 | 0.97895 | 1.034 | 0.95531 | 0.0092398 | -0.014529 | 0.019855 | 79.6078 | -2 | -4 | -0.39216 | 1 | -2 | 2.2702 | 2.2361 | zero_calibration | 0
A | 5 | A5 | 1 | C | 4 | 1 | 0.37868 | 0.30663 | 0.56681 | 0.53275 | 0.56064 | 0.63705 | 0.71079 | 0.54694 | 0.88975 | 0.14826 | 0.26206 | 0.050733 | 64.7059 | 15 | -22 | -15.2941 | 18 | -20 | 30.9501 | 26.9072 | zero_calibration | 0
A | 6 | A6 | 1 | C | 4 | 1 | 0.37362 | 0.29765 | 0.5605 | 0.53128 | 0.5605 | 0.63245 | 0.70323 | 0.53105 | 0.88623 | 0.1529 | 0.27487 | 0.052452 | 63.9216 | 16 | -23 | -16.0784 | 19 | -21 | 32.5656 | 28.3196 | zero_calibration | 0
A | 7 | A7 | 1 | C | 4 | 1 | 0.37868 | 0.29322 | 0.5605 | 0.53029 | 0.56006 | 0.62693 | 0.7141 | 0.52354 | 0.89404 | 0.14624 | 0.28105 | 0.048641 | 63.9216 | 17 | -23 | -16.0784 | 20 | -21 | 33.1589 | 29 | zero_calibration | 0
A | 8 | A8 | 1 | C | 4 | 1 | 0.38377 | 0.30213 | 0.55423 | 0.52976 | 0.55932 | 0.62048 | 0.72444 | 0.54017 | 0.89322 | 0.14 | 0.26747 | 0.049042 | 64.3137 | 15 | -22 | -15.6863 | 18 | -20 | 31.1458 | 26.9072 | zero_calibration | 0
B | 1 | B1 | 1 | C | 0.5 | 1 | 0.52952 | 0.55423 | 0.61207 | 0.55258 | 0.5705 | 0.65916 | 0.95828 | 0.97148 | 0.92856 | 0.018508 | 0.012567 | 0.03219 | 78.8235 | 0 | -5 | -1.1765 | 3 | -3 | 4.4027 | 4.2426 | zero_calibration | 0
B | 2 | B2 | 1 | C | 0.5 | 1 | 0.5418 | 0.56681 | 0.63878 | 0.54943 | 0.57273 | 0.65814 | 0.98611 | 0.98967 | 0.97059 | 0.0060757 | 0.0045094 | 0.012965 | 79.6078 | 1 | -6 | -0.39216 | 4 | -4 | 5.6704 | 5.6569 | zero_calibration | 0
B | 3 | B3 | 1 | C | 0.5 | 1 | 0.53564 | 0.56681 | 0.63878 | 0.54677 | 0.57465 | 0.65614 | 0.97964 | 0.98635 | 0.97354 | 0.0089332 | 0.0059683 | 0.011646 | 79.6078 | 1 | -6 | -0.39216 | 4 | -4 | 5.6704 | 5.6569 | zero_calibration | 0
B | 4 | B4 | 1 | C | 0.5 | 1 | 0.52952 | 0.55423 | 0.63878 | 0.5446 | 0.57628 | 0.65317 | 0.97231 | 0.96173 | 0.97796 | 0.012193 | 0.016945 | 0.0096781 | 78.8235 | 1 | -6 | -1.1765 | 4 | -4 | 5.7779 | 5.6569 | zero_calibration | 0
B | 5 | B5 | 1 | C | 4.5 | 1 | 0.38377 | 0.29765 | 0.59244 | 0.54291 | 0.5776 | 0.64924 | 0.70688 | 0.51533 | 0.9125 | 0.15065 | 0.28792 | 0.039765 | 64.3137 | 18 | -26 | -15.6863 | 21 | -24 | 35.5395 | 31.8904 | zero_calibration | 0
B | 6 | B6 | 1 | C | 4.5 | 1 | 0.37362 | 0.28445 | 0.58597 | 0.5417 | 0.57862 | 0.64436 | 0.68971 | 0.4916 | 0.90939 | 0.16133 | 0.30838 | 0.041252 | 63.1373 | 19 | -26 | -16.8627 | 22 | -24 | 36.6654 | 32.5576 | zero_calibration | 0
B | 7 | B7 | 1 | C | 4.5 | 1 | 0.37362 | 0.27583 | 0.57955 | 0.54096 | 0.57933 | 0.63854 | 0.69065 | 0.47612 | 0.90762 | 0.16074 | 0.32228 | 0.042097 | 62.7451 | 19 | -27 | -17.2549 | 22 | -25 | 37.5064 | 33.3017 | zero_calibration | 0
B | 8 | B8 | 1 | C | 4.5 | 1 | 0.38377 | 0.28882 | 0.57955 | 0.54071 | 0.57974 | 0.63178 | 0.70976 | 0.49818 | 0.91732 | 0.14889 | 0.30261 | 0.03748 | 63.5294 | 18 | -25 | -16.4706 | 21 | -23 | 35.2318 | 31.1448 | zero_calibration | 0
C | 1 | C1 | 1 | C | 1 | 1 | 0.5174 | 0.52952 | 0.62534 | 0.55894 | 0.57827 | 0.66848 | 0.92569 | 0.9157 | 0.93547 | 0.033536 | 0.038248 | 0.028971 | 77.6471 | 2 | -8 | -2.3529 | 5 | -6 | 8.157 | 7.8102 | zero_calibration | 0
C | 2 | C2 | 1 | C | 1 | 1 | 0.52952 | 0.53564 | 0.65237 | 0.55604 | 0.58167 | 0.66721 | 0.95231 | 0.92087 | 0.97775 | 0.021219 | 0.035802 | 0.0097709 | 78.0392 | 3 | -9 | -1.9608 | 6 | -7 | 9.4257 | 9.2195 | zero_calibration | 0
C | 3 | C3 | 1 | C | 1 | 1 | 0.52344 | 0.52952 | 0.65922 | 0.55363 | 0.58477 | 0.66496 | 0.94547 | 0.90552 | 0.99138 | 0.02435 | 0.0431 | 0.0037616 | 78.0392 | 4 | -10 | -1.9608 | 7 | -8 | 10.8095 | 10.6301 | zero_calibration | 0
C | 4 | C4 | 1 | C | 1 | 1 | 0.5174 | 0.52952 | 0.65922 | 0.55171 | 0.58757 | 0.66172 | 0.93781 | 0.90121 | 0.99622 | 0.027884 | 0.045176 | 0.0016437 | 78.0392 | 3 | -10 | -1.9608 | 6 | -8 | 10.1904 | 10 | zero_calibration | 0
C | 5 | C5 | 1 | C | 5 | 1 | 0.38377 | 0.29322 | 0.59894 | 0.55028 | 0.59007 | 0.65752 | 0.69742 | 0.49691 | 0.91091 | 0.15651 | 0.30372 | 0.040522 | 63.9216 | 18 | -26 | -16.0784 | 21 | -24 | 35.7144 | 31.8904 | zero_calibration | 0
C | 6 | C6 | 1 | C | 5 | 1 | 0.37868 | 0.28882 | 0.59244 | 0.54933 | 0.59227 | 0.65235 | 0.68934 | 0.48764 | 0.90816 | 0.16156 | 0.3119 | 0.041837 | 63.5294 | 18 | -26 | -16.4706 | 21 | -24 | 35.8926 | 31.8904 | zero_calibration | 0
C | 7 | C7 | 1 | C | 5 | 1 | 0.37868 | 0.28012 | 0.59244 | 0.54886 | 0.59417 | 0.64623 | 0.68993 | 0.47146 | 0.91676 | 0.16119 | 0.32656 | 0.037745 | 63.1373 | 19 | -27 | -16.8627 | 22 | -25 | 37.3276 | 33.3017 | zero_calibration | 0
C | 8 | C8 | 1 | C | 5 | 1 | 0.38377 | 0.28445 | 0.58597 | 0.54888 | 0.59576 | 0.63918 | 0.6992 | 0.47746 | 0.91676 | 0.1554 | 0.32106 | 0.037746 | 63.5294 | 19 | -26 | -16.4706 | 22 | -24 | 36.4867 | 32.5576 | zero_calibration | 0
D | 1 | D1 | 1 | C | 1.5 | 1 | 0.49362 | 0.48195 | 0.63204 | 0.56243 | 0.58141 | 0.67381 | 0.87765 | 0.82894 | 0.93801 | 0.056681 | 0.081476 | 0.027791 | 75.2941 | 5 | -12 | -4.7059 | 8 | -10 | 13.6435 | 12.8062 | zero_calibration | 0
D | 2 | D2 | 1 | C | 1.5 | 1 | 0.50543 | 0.49362 | 0.65922 | 0.55978 | 0.58596 | 0.67228 | 0.90291 | 0.8424 | 0.98058 | 0.044357 | 0.074481 | 0.0085192 | 76.0784 | 6 | -13 | -3.9216 | 9 | -11 | 14.7438 | 14.2127 | zero_calibration | 0
D | 3 | D3 | 1 | C | 1.5 | 1 | 0.50543 | 0.48777 | 0.65922 | 0.55763 | 0.59023 | 0.66976 | 0.90639 | 0.8264 | 0.98426 | 0.042683 | 0.082812 | 0.0068892 | 75.6863 | 6 | -13 | -4.3137 | 9 | -11 | 14.8529 | 14.2127 | zero_calibration | 0
D | 4 | D4 | 1 | C | 1.5 | 1 | 0.50543 | 0.49362 | 0.65922 | 0.55597 | 0.59421 | 0.66626 | 0.9091 | 0.83071 | 0.98944 | 0.041386 | 0.080552 | 0.0046112 | 76.0784 | 6 | -13 | -3.9216 | 9 | -11 | 14.7438 | 14.2127 | zero_calibration | 0
E | 1 | E1 | 1 | C | 2 | 1 | 0.47044 | 0.43681 | 0.62534 | 0.56303 | 0.57986 | 0.67508 | 0.83555 | 0.75331 | 0.92632 | 0.078026 | 0.12302 | 0.033238 | 72.9412 | 8 | -15 | -7.0588 | 11 | -13 | 18.4344 | 17.0294 | zero_calibration | 0
E | 2 | E2 | 1 | C | 2 | 1 | 0.48195 | 0.44787 | 0.65237 | 0.56064 | 0.58555 | 0.67329 | 0.85965 | 0.76487 | 0.96892 | 0.065677 | 0.11641 | 0.013711 | 73.7255 | 9 | -16 | -6.2745 | 12 | -14 | 19.4774 | 18.4391 | zero_calibration | 0
E | 3 | E3 | 1 | C | 2 | 1 | 0.48195 | 0.44787 | 0.65237 | 0.55874 | 0.59097 | 0.67051 | 0.86257 | 0.75786 | 0.97295 | 0.064205 | 0.12041 | 0.011912 | 73.3333 | 9 | -16 | -6.6667 | 12 | -14 | 19.6073 | 18.4391 | zero_calibration | 0
E | 4 | E4 | 1 | C | 2 | 1 | 0.48777 | 0.45908 | 0.65237 | 0.55734 | 0.59611 | 0.66674 | 0.87517 | 0.77013 | 0.97845 | 0.057906 | 0.11344 | 0.0094621 | 74.5098 | 8 | -15 | -5.4902 | 11 | -13 | 17.8925 | 17.0294 | zero_calibration | 0
F | 1 | F1 | 1 | C | 2.5 | 1 | 0.45346 | 0.40454 | 0.60548 | 0.56072 | 0.57364 | 0.67229 | 0.8087 | 0.70521 | 0.90063 | 0.092214 | 0.15168 | 0.045453 | 70.9804 | 11 | -16 | -9.0196 | 14 | -14 | 21.7567 | 19.799 | zero_calibration | 0
F | 2 | F2 | 1 | C | 2.5 | 1 | 0.45908 | 0.40983 | 0.63878 | 0.55859 | 0.58044 | 0.67024 | 0.82186 | 0.70607 | 0.95307 | 0.085204 | 0.15116 | 0.020877 | 71.3726 | 11 | -18 | -8.6274 | 14 | -16 | 22.9441 | 21.2603 | zero_calibration | 0
F | 3 | F3 | 1 | C | 2.5 | 1 | 0.45908 | 0.40454 | 0.63878 | 0.55695 | 0.58697 | 0.66719 | 0.82428 | 0.6892 | 0.95742 | 0.083927 | 0.16165 | 0.018898 | 71.3726 | 12 | -19 | -8.6274 | 15 | -17 | 24.2576 | 22.6716 | zero_calibration | 0
F | 4 | F4 | 1 | C | 2.5 | 1 | 0.46474 | 0.41515 | 0.63878 | 0.5558 | 0.59324 | 0.66316 | 0.83616 | 0.6998 | 0.96324 | 0.077709 | 0.15503 | 0.016264 | 72.1569 | 10 | -17 | -7.8431 | 13 | -15 | 21.3428 | 19.8494 | zero_calibration | 0
G | 1 | G1 | 1 | C | 3 | 1 | 0.43681 | 0.37362 | 0.60548 | 0.55554 | 0.56285 | 0.66546 | 0.78629 | 0.66379 | 0.90988 | 0.10442 | 0.17797 | 0.041018 | 69.4118 | 12 | -19 | -10.5882 | 15 | -17 | 25.0222 | 22.6716 | zero_calibration | 0
G | 2 | G2 | 1 | C | 3 | 1 | 0.44232 | 0.37362 | 0.62534 | 0.55366 | 0.57069 | 0.66314 | 0.79891 | 0.65467 | 0.943 | 0.097503 | 0.18397 | 0.025486 | 69.4118 | 13 | -20 | -10.5882 | 16 | -18 | 26.308 | 24.0832 | zero_calibration | 0
G | 3 | G3 | 1 | C | 3 | 1 | 0.44232 | 0.37362 | 0.62534 | 0.55227 | 0.57829 | 0.65983 | 0.80091 | 0.64607 | 0.94773 | 0.096415 | 0.18972 | 0.023315 | 69.4118 | 13 | -20 | -10.5882 | 16 | -18 | 26.308 | 24.0832 | zero_calibration | 0
G | 4 | G4 | 1 | C | 3 | 1 | 0.44787 | 0.38891 | 0.62534 | 0.55138 | 0.58564 | 0.65554 | 0.81227 | 0.66407 | 0.95393 | 0.090298 | 0.17778 | 0.020483 | 70.1961 | 12 | -19 | -9.8039 | 15 | -17 | 24.7005 | 22.6716 | zero_calibration | 0
H | 1 | H1 | 1 | C | 3.5 | 1 | 0.41515 | 0.34886 | 0.59244 | 0.54751 | 0.54762 | 0.65466 | 0.75824 | 0.63706 | 0.90496 | 0.12019 | 0.19582 | 0.04337 | 67.8431 | 13 | -20 | -12.1569 | 16 | -18 | 26.9776 | 24.0832 | zero_calibration | 0
H | 2 | H2 | 1 | C | 3.5 | 1 | 0.41515 | 0.33922 | 0.59244 | 0.54589 | 0.55644 | 0.65208 | 0.7605 | 0.60963 | 0.90854 | 0.1189 | 0.21493 | 0.041656 | 67.0588 | 15 | -21 | -12.9412 | 18 | -19 | 29.1972 | 26.1725 | zero_calibration | 0
H | 3 | H3 | 1 | C | 3.5 | 1 | 0.42051 | 0.34403 | 0.59244 | 0.54475 | 0.56504 | 0.64852 | 0.77192 | 0.60885 | 0.91353 | 0.11243 | 0.21549 | 0.039279 | 67.0588 | 15 | -21 | -12.9412 | 18 | -19 | 29.1972 | 26.1725 | zero_calibration | 0
H | 4 | H4 | 1 | C | 3.5 | 1 | 0.42051 | 0.35374 | 0.59894 | 0.54411 | 0.57342 | 0.64398 | 0.77284 | 0.6169 | 0.93006 | 0.11191 | 0.20978 | 0.03149 | 67.8431 | 13 | -20 | -12.1569 | 16 | -18 | 26.9776 | 24.0832 | zero_calibration | 0

## 05_REPLICATES_MEAN
ID | DF | Type | Conc | PAbs_Red_median | PAbs_Red_sd | PAbs_Green_median | PAbs_Green_sd | PAbs_Blue_median | PAbs_Blue_sd | L_median | L_sd | a_median | a_sd | b_median | b_sd | DeltaL_median | DeltaL_sd | Deltaa_median | Deltaa_sd | Deltab_median | Deltab_sd | DeltaE_ab_median | DeltaE_ab_sd | DeltaE_ab_chroma_median | DeltaE_ab_chroma_sd | CIELAB_ref_source | NReplicates | QCFlagged | QCCritical
1 | 1 | C | 0 | 0.010221 | 0.0019843 | -0.020388 | 0.0039727 | 0.026964 | 0.0047573 | 80 | 0.29071 | -3 | 0 | -2 | 0 | 0 | 0.29071 | 0 | 0 | 0 | 0 | 0.19608 | 0.29071 | 0 | 0 | zero_calibration | 4 | 0 | 0
1 | 1 | C | 0.5 | 0.010563 | 0.004535 | 0.0092676 | 0.005973 | 0.012305 | 0.0024364 | 79.2157 | 0.58141 | 1 | 0 | -6 | 0 | -0.78431 | 0.58141 | 4 | 0 | -4 | 0 | 5.6704 | 0.079663 | 5.6569 | 0 | zero_calibration | 4 | 0 | 0
1 | 1 | C | 1 | 0.026117 | 0.0049403 | 0.040674 | 0.0051358 | 0.0067662 | 0.0060246 | 78.0392 | 0 | 3 | 0.7413 | -9.5 | 0.7413 | -1.9608 | 0 | 6 | 0.7413 | -7.5 | 0.7413 | 9.8081 | 1.0258 | 9.6098 | 1.0457 | zero_calibration | 4 | 0 | 0
1 | 1 | C | 1.5 | 0.04352 | 0.0022022 | 0.081014 | 0.0016757 | 0.0077042 | 0.002897 | 75.8824 | 0.2907 | 6 | 0 | -13 | 0 | -4.1176 | 0.2907 | 9 | 0 | -11 | 0 | 14.7438 | 0.080888 | 14.2127 | 0 | zero_calibration | 4 | 0 | 0
1 | 1 | C | 2 | 0.064941 | 0.0057612 | 0.11841 | 0.0049027 | 0.012811 | 0.0031495 | 73.5294 | 0.58141 | 8.5 | 0.7413 | -15.5 | 0.7413 | -6.4706 | 0.58141 | 11.5 | 0.7413 | -13.5 | 0.7413 | 18.9559 | 0.86944 | 17.7342 | 1.045 | zero_calibration | 4 | 0 | 0
1 | 1 | C | 2.5 | 0.084566 | 0.0055565 | 0.15335 | 0.0028705 | 0.019888 | 0.0034195 | 71.3726 | 0.29071 | 11 | 0.7413 | -17.5 | 1.4826 | -8.6274 | 0.29071 | 14 | 0.7413 | -15.5 | 1.4826 | 22.3504 | 1.1871 | 20.5549 | 1.0833 | zero_calibration | 4 | 0 | 0
1 | 1 | C | 3 | 0.096959 | 0.0053409 | 0.18097 | 0.0045901 | 0.024401 | 0.0037093 | 69.4118 | 0 | 12.5 | 0.7413 | -19.5 | 0.7413 | -10.5882 | 0 | 15.5 | 0.7413 | -17.5 | 0.7413 | 25.6651 | 0.95316 | 23.3774 | 1.0464 | zero_calibration | 4 | 0 | 0
1 | 1 | C | 3.5 | 0.11566 | 0.0051813 | 0.21236 | 0.0042276 | 0.040468 | 0.0030326 | 67.451 | 0.58141 | 14 | 1.4826 | -20.5 | 0.7413 | -12.549 | 0.58141 | 17 | 1.4826 | -18.5 | 0.7413 | 28.0874 | 1.6454 | 25.1278 | 1.5488 | zero_calibration | 4 | 0 | 0
1 | 1 | C | 4 | 0.14725 | 0.0049378 | 0.27117 | 0.009491 | 0.049888 | 0.001551 | 64.1176 | 0.29071 | 15.5 | 0.7413 | -22.5 | 0.7413 | -15.8824 | 0.29071 | 18.5 | 0.7413 | -20.5 | 0.7413 | 31.8557 | 1.1975 | 27.6134 | 1.047 | zero_calibration | 4 | 0 | 0
1 | 1 | C | 4.5 | 0.1557 | 0.0079184 | 0.3055 | 0.014581 | 0.040509 | 0.0017285 | 63.3333 | 0.58141 | 18.5 | 0.7413 | -26 | 0.7413 | -16.6667 | 0.58141 | 21.5 | 0.7413 | -24 | 0.7413 | 36.1025 | 1.0627 | 32.224 | 1.0461 | zero_calibration | 4 | 0 | 0
1 | 1 | C | 5 | 0.15885 | 0.0037498 | 0.31648 | 0.010866 | 0.039134 | 0.0020588 | 63.5294 | 0.2907 | 18.5 | 0.7413 | -26 | 0 | -16.4706 | 0.2907 | 21.5 | 0.7413 | -24 | 0 | 36.1897 | 0.57254 | 32.224 | 0.4946 | zero_calibration | 4 | 0 | 0

## 06_FITTING
Channel | FitType | n_points | m | q | R2 | RMSE | sigma_cal | sigma_source | SNR | LOD | LOQ | S0_calibration | S0_applied | NClipPoints | ClipX | ClipDelta
PAbs_Red | Calibration | 11 | 0.029967 | 0 | 0.98393 | 0.0063735 | 0.0049403 | median_calibration_sd | 6.0657 | 0.49458 | 1.6486 | 0 | 0 | 3 | 0,0.5,1,1.5,2,2.5,3,3.5,4,4.5,5 | 0,0.0147,0.014189,0.011828,0,0,0,0,0,0,0
PAbs_Green | Calibration | 11 | 0.068738 | 0 | 0.99525 | 0.0076953 | 0.0049027 | median_calibration_sd | 14.0203 | 0.21397 | 0.71325 | 0 | 0 | 2 | 0,0.5,1,1.5,2,2.5,3,3.5,4,4.5,5 | 0,0,0.0074425,0,0,0,0,0.0070214,0,0,0
PAbs_Blue | Calibration | 11 | 0.0073746 | 0 | 0.86643 | 0.0049988 | 0.0030326 | median_calibration_sd | 2.4318 | 1.2337 | 4.1122 | 0.020198 | 0.020198 | 5 | 0,0.5,1,1.5,2,2.5,3,3.5,4,4.5,5 | 0,0,0.0070894,0.0096961,0.0081338,0.004602,0,0,0,0,0.0030792
L | Calibration | 11 | -3.9186 | 1.2237 | 0.97997 | 0.8569 | 0.29071 | median_calibration_sd | 13.4795 | 0.22256 | 0.74187
a | Calibration | 11 | 4.2507 | 2.072 | 0.97759 | 1.0227 | 0.7413 | median_calibration_sd | 5.7341 | 0.52318 | 1.7439
b | Calibration | 11 | -4.6445 | -2.6171 | 0.96941 | 1.3174 | 0.7413 | median_calibration_sd | 6.2653 | 0.47883 | 1.5961
DeltaL | Calibration | 11 | -3.9186 | 1.2237 | 0.97997 | 0.8569 | 0.29071 | median_calibration_sd | 13.4795 | 0.22256 | 0.74187
Deltaa | Calibration | 11 | 4.2507 | 2.072 | 0.97759 | 1.0227 | 0.7413 | median_calibration_sd | 5.7341 | 0.52318 | 1.7439
Deltab | Calibration | 11 | -4.6445 | -2.6171 | 0.96941 | 1.3174 | 0.7413 | median_calibration_sd | 6.2653 | 0.47883 | 1.5961
DeltaE | Calibration | 11 | 7.3373 | 2.3778 | 0.98439 | 1.4524 | 0.95316 | median_calibration_sd | 7.6979 | 0.38972 | 1.2991
DeltaE_chroma | Calibration | 11 | 6.284 | 3.2887 | 0.97495 | 1.6075 | 1.0457 | median_calibration_sd | 6.0095 | 0.49921 | 1.664

## 07_METHOD_COMPARISON
Selected | Rank | Method | Family | ComparableGroup | CommonFactorsN | Score | ScoreFormula | RankMode | R2_cal | m_cal | SNR | LOD | LOQ | n_stdadd | n_unknown | BaseScore
1 | 1 | PAbs_Green | RGB | calibration_only | 1 | 0.99525 | R2_cal | calibration_only | 0.99525 | 0.068738 | 14.0203 | 0.21397 | 0.71325 | 0 | 0 | 0.99525
0 | 2 | DeltaE | DeltaCIELAB | calibration_only | 1 | 0.98439 | R2_cal | calibration_only | 0.98439 | 7.3373 | 7.6979 | 0.38972 | 1.2991 | 0 | 0 | 0.98439
0 | 3 | PAbs_Red | RGB | calibration_only | 1 | 0.98393 | R2_cal | calibration_only | 0.98393 | 0.029967 | 6.0657 | 0.49458 | 1.6486 | 0 | 0 | 0.98393
0 | 4 | L | CIELAB | calibration_only | 1 | 0.97997 | R2_cal | calibration_only | 0.97997 | -3.9186 | 13.4795 | 0.22256 | 0.74187 | 0 | 0 | 0.97997
0 | 5 | DeltaL | DeltaCIELAB | calibration_only | 1 | 0.97997 | R2_cal | calibration_only | 0.97997 | -3.9186 | 13.4795 | 0.22256 | 0.74187 | 0 | 0 | 0.97997
0 | 6 | a | CIELAB | calibration_only | 1 | 0.97759 | R2_cal | calibration_only | 0.97759 | 4.2507 | 5.7341 | 0.52318 | 1.7439 | 0 | 0 | 0.97759
0 | 7 | Deltaa | DeltaCIELAB | calibration_only | 1 | 0.97759 | R2_cal | calibration_only | 0.97759 | 4.2507 | 5.7341 | 0.52318 | 1.7439 | 0 | 0 | 0.97759
0 | 8 | DeltaE_chroma | DeltaCIELAB | calibration_only | 1 | 0.97495 | R2_cal | calibration_only | 0.97495 | 6.284 | 6.0095 | 0.49921 | 1.664 | 0 | 0 | 0.97495
0 | 9 | b | CIELAB | calibration_only | 1 | 0.96941 | R2_cal | calibration_only | 0.96941 | -4.6445 | 6.2653 | 0.47883 | 1.5961 | 0 | 0 | 0.96941
0 | 10 | Deltab | DeltaCIELAB | calibration_only | 1 | 0.96941 | R2_cal | calibration_only | 0.96941 | -4.6445 | 6.2653 | 0.47883 | 1.5961 | 0 | 0 | 0.96941
0 | 11 | PAbs_Blue | RGB | calibration_only | 1 | 0.86643 | R2_cal | calibration_only | 0.86643 | 0.0073746 | 2.4318 | 1.2337 | 4.1122 | 0 | 0 | 0.86643

## 08_LEGENDS
Term | Meaning | Formula | Unit | Where used | Shown when | Notes
a | CIELAB green–red axis | median ROI value after RGB → CIELAB conversion | dimensionless | RAW, SUMMARY | CIELAB enabled
b | CIELAB blue–yellow axis | median ROI value after RGB → CIELAB conversion | dimensionless | RAW, SUMMARY | CIELAB enabled
BaseScore/FinalScore | Intermediate/final score fields when present | same common score unless an explicit post-score adjustment is applied | dimensionless | METHOD_COMPARISON | method comparison | If identical to Score in the exported workbook, Score is the authoritative field.
beta_k | Slope ratio for one standard-addition curve | m_std / m_cal | dimensionless | FITTING, METHOD_COMPARISON | stdadd present | Used to assess whether StdAdd and calibration slopes are coherent.
beta_k/beta_mean | Standard-addition/calibration slope ratio | m_std / m_cal; beta_mean is the mean across available standard-addition fits | dimensionless | 06_FITTING, 07_METHOD_COMPARISON | calibration plus standard addition present
bias_index_k | Absolute relative slope bias for one standard-addition curve | ¦beta_k − 1¦ | dimensionless | FITTING, METHOD_COMPARISON | stdadd present | 0 means identical slopes.
bias_index_k/bias_index_mean | Relative slope-bias index | ¦m_std/m_cal - 1¦; bias_index_mean is the mean across available fits | dimensionless | 06_FITTING, 07_METHOD_COMPARISON | calibration plus standard addition present
BiasIndex | Relative slope bias between standard addition and calibration | ¦β − 1¦, where β = m_std / m_cal | dimensionless | METHOD_COMPARISON | stdadd present | Equals 0 when the standard-addition slope equals the calibration slope. Larger values indicate stronger slope bias.
C0 | Concentration in original sample | standard addition: DF × (q/m); unknown/CRM from calibration: DF × ((y − q)/m); unknown from epsilon: DF × PAbs/(epsilon × l_cm), converted from M to the selected unit | mM | FITTING, OVERVIEW | stdadd present or unknown/CRM present | Same symbol is used for standard addition and unknown/CRM results; FitType defines the method.
C0/C0_sd | Estimated original-sample concentration and associated uncertainty | standard addition: DF x q/m; calibration projection: DF x (y - q)/m; epsilon mode: DF x PAbs/(epsilon x l_cm), converted from M to the selected unit | mM | FIGURE_RGB.png, 03_OVERVIEW, 06_FITTING | standard addition, unknown or epsilon quantification present | FitType identifies which calculation path produced C0.
C0_sd | Uncertainty on C0 | Fit covariance for calibration/std-add; for epsilon mode, propagated PAbs replicate SD divided by epsilon × l_cm and converted to the selected unit | mM | FITTING, OVERVIEW | stdadd present or unknown/CRM present | Reported in the same concentration units as C0.
Channel | Fitted analytical descriptor | PAbs_Red, PAbs_Green, PAbs_Blue or diagnostic CIELAB/DeltaE descriptor | text | FITTING, METHOD_COMPARISON | always
CIELAB (L*, a*, b*) | CIE 1976 color coordinates derived from the RGB image | L* = 116 f(Y/Yn) - 16; a* = 500[f(X/Xn) - f(Y/Yn)]; b* = 200[f(Y/Yn) - f(Z/Zn)] | dimensionless | RAW, REPLICATES_MEAN, CIELAB diagnostics | CIELAB enabled | D65 is the reference white used by the sRGB/OpenCV conversion convention. References: CIE 1976 L*a*b* and CIE standard illuminant D65.
CIELAB_ref_source | Source of the CIELAB reference used to compute Delta variables | zero calibration or other reported fallback source | text | 04_RAW, 05_REPLICATES_MEAN | CIELAB/DeltaE outputs present | Reported only for CIELAB-derived variables because RGB pseudo-absorbance does not use this CIELAB reference.
ClipDelta | Positive calibration-derived clipping correction | max(0, y_expected − y_shifted) when deficit exceeds SD threshold | signal | FITTING | RGB calibration | Computed from calibration after S0 and zero-intercept fit; applied by same channel and same x.
ClipX | Calibration x values where clipping correction was evaluated | list of calibration concentrations | mM | FITTING | RGB calibration | Transferred to StdAdd at the same channel and same x values.
Code implementation | Main computational libraries and code provenance | custom Python code using NumPy, OpenCV, openpyxl and matplotlib | software provenance | all generated outputs | always | Sources/libraries: NumPy for least-squares solves; OpenCV for image masks/color conversion; openpyxl for xlsx export; matplotlib for PNG figures.
Col | Plate column index | 1-based column number | well-column index | RAW, REPLICATES_MEAN | always
CommonFactorsN | Number of common factors used in the score | 3 for R²_cal, R²_std, slope agreement; 1 for calibration-only or stdadd-only fallbacks | integer | METHOD_COMPARISON | always | Used to avoid comparing scores obtained from different formulas.
ComparableGroup | Set of methods scored with the same formula | calibration_plus_stdadd, calibration_only, stdadd_only, or not_ranked | text | METHOD_COMPARISON | always | Scores are directly comparable only within the same group. The group with the largest CommonFactorsN is used for Selected/Rank.
Conc | Configured concentration for calibration or added-standard wells | user/configurator input | mM | RAW, REPLICATES_MEAN, FITTING | always | Unknown wells may have undefined Conc.
Confidence class | Qualitative class derived from reliability score | HIGH >= 75; MEDIUM >= 45; LOW < 45; NOT QUANTIFIABLE when no valid quantification is available | class | OVERVIEW | always | Interpret together with the reliability reason.
delta_expected_*/delta_reference_* | Difference between estimate and external reference | estimate - reference | mM | METHOD_COMPARISON.png, 03_OVERVIEW, 07_METHOD_COMPARISON | reference values configured
delta_reference_* | Difference between estimated concentration and an external reference | C0 - reference_value | mM | OVERVIEW, METHOD_COMPARISON, FIGURE_RGB | reference values configured
Deltaa | Difference in CIELAB a* relative to the selected reference | a - a_ref | dimensionless | RAW, SUMMARY | CIELAB enabled | Reference source is reported in CIELAB_ref_source.
Deltab | Difference in CIELAB b* relative to the selected reference | b - b_ref | dimensionless | RAW, SUMMARY | CIELAB enabled | Reference source is reported in CIELAB_ref_source.
DeltaE_ab | Total CIELAB color difference | sqrt(DeltaL^2 + Deltaa^2 + Deltab^2) | dimensionless | FIGURE_CIELAB_DELTAE.png, RAW_DATA_DETAILS_CAPTION.txt, 04_RAW, 05_REPLICATES_MEAN, 07_METHOD_COMPARISON, diagnostics | DeltaE outputs present | Reference: CIE 1976 L*a*b* color-difference form.
DeltaE_ab_chroma | Chromatic CIELAB color difference without lightness | sqrt(Deltaa^2 + Deltab^2) | dimensionless | FIGURE_CIELAB_DELTAE.png, RAW_DATA_DETAILS_CAPTION.txt, 04_RAW, 05_REPLICATES_MEAN, diagnostics | DeltaE outputs present | Reference source is reported in CIELAB_ref_source.
DeltaL | Difference in CIELAB lightness relative to the selected reference | L - L_ref | dimensionless | RAW, SUMMARY | CIELAB enabled | Reference source is reported in CIELAB_ref_source.
DeltaL/Deltaa/Deltab | CIELAB coordinate differences from the selected CIELAB reference | DeltaL = L - L_ref; Deltaa = a - a_ref; Deltab = b - b_ref | dimensionless | FIGURE_CIELAB_DELTAE.png, 04_RAW, 05_REPLICATES_MEAN, diagnostics | Delta outputs present | The reference source is reported in CIELAB_ref_source.
DF | Dilution factor used to report original-sample concentration | C0 = DF x concentration in analyzed/diluted well | dimensionless | RAW, REPLICATES_MEAN, FITTING | always
Empty-well QC | QC based on empty-well robust SD and/or stored empty comparison | warning/watch/ok/not_available from drift score and robust SD thresholds | class | OVERVIEW | always | Flags comparability or background/empty-well drift.
epsilon | User-configured Beer-Lambert-like proportionality coefficient for PAbs quantification | PAbs = epsilon × l_cm × C_M | M-1 cm-1 | FITTING, OVERVIEW | unknown-only epsilon mode | C_M is mol/L. The analyzer converts the calculated concentration to the selected output unit.
Estimate_value/Estimate_sd | Representative concentration estimate and associated SD used for method comparison | derived from standard addition, calibration projection, or epsilon mode depending on Estimate_source | mM | METHOD_COMPARISON | method comparison
Estimate_value/Estimate_sd/Estimate_source | Representative concentration estimate, associated SD and source | selected according to Estimate_source | mM and text | METHOD_COMPARISON.png, 07_METHOD_COMPARISON | method comparison present
expected_*/reference_* | External reference metadata configured by the user | user/configurator input | text or mM | METHOD_COMPARISON.png, 03_OVERVIEW, 07_METHOD_COMPARISON | reference values configured | External reference values are checks only and are not used for ranking.
Family | Method family used in method comparison | RGB, CIELAB, DeltaCIELAB, or other | text | METHOD_COMPARISON | always
Field | Name of a metadata or overview item | label reported in the first column of the sheet | text | 02_METADATA, 03_OVERVIEW | always
FitType | Type of fit or concentration-estimation row | Calibration, StdAdd, UnknownFromCal, UnknownFromEpsilon, UnknownOnly | text | FITTING | always
flatfield_span | Image-level slow-field nonuniformity | (P95 − P5) / median of the slow grayscale field | dimensionless | METADATA | always | Part of rule-based image QC; larger values indicate stronger illumination/background gradient.
FracWellWarnings/FracWellCritical | Fractions of replicate wells with warning or critical optical QC | count / NReplicates | dimensionless | REPLICATES_MEAN | always
Geometry and epsilon/path-length quantification | Assumption used when epsilon-based unknown quantification is enabled | l_cm = (volume_uL / well_bottom_area_mm2) / 10; C_M = PAbs / (epsilon x l_cm) | cm and M | RESULTS_CAPTION.txt, RAW_DATA_DETAILS_CAPTION.txt, REPORT | epsilon/path-length mode configured | Assumes ANSI/SLAS-compatible flat-bottom microplate geometry; non-flat or non-certified geometries require separate validation.
ID | Sample or reference identifier assigned in the plate map | user/configurator input | text | RAW, REPLICATES_MEAN, FITTING, METHOD_COMPARISON | always
ImageWarning | Well-level optical QC warning flag | 1 if optical QC rules flag the well, else 0 | 0/1 | 04_RAW | always
ImageWarning_any | Group-level indicator that at least one replicate had an image warning | 1 if any replicate ImageWarning = 1 | 0/1 | REPLICATES_MEAN | always
IRLS | Iteratively reweighted least-squares robust linear regression with residual-based weights | minimize sum_i w_i (y_i - (m x_i + q))^2; w_i is updated iteratively from residual magnitude | dimensionless | FIGURE_RGB.png, FIGURE_CIELAB_DELTAE.png, RESULTS_CAPTION.txt, RAW_DATA_DETAILS_CAPTION.txt, 06_FITTING, 07_METHOD_COMPARISON, diagnostic fitting tables | always | Reference: Huber, P. J. (1964), Robust Estimation of a Location Parameter. Implementation: custom NumPy-based IRLS using repeated least-squares solves.
L | CIELAB lightness | median ROI value after RGB → CIELAB conversion | dimensionless | RAW, SUMMARY | CIELAB enabled
L/a/b | CIELAB lightness and opponent-color coordinates | median ROI value after RGB to CIELAB conversion | dimensionless | FIGURE_CIELAB_DELTAE.png, 04_RAW, 05_REPLICATES_MEAN, diagnostics | CIELAB outputs present
liquid_volume_ul | Liquid volume loaded into each well | user input | uL | OVERVIEW, FITTING | epsilon mode configured | Used only to calculate path length for flat-bottom wells.
LOD | Limit of detection | 3 × σ / ¦m¦ | mM | FITTING, OVERVIEW | calibration only or calibration present | Shown only for calibration rows.
LOD/LOQ | Detection and quantification limits from calibration | LOD = 3 x sigma_cal / ¦m¦; LOQ = 10 x sigma_cal / ¦m¦ | mM | FIGURE_RGB.png, FIGURE_CIELAB_DELTAE.png, 03_OVERVIEW, 06_FITTING, 07_METHOD_COMPARISON | calibration present
LOQ | Limit of quantification | 10 × σ / ¦m¦ | mM | FITTING, OVERVIEW | calibration only or calibration present | Shown only for calibration rows.
m | Slope of linear fit | y = m x + q | signal / mM | FITTING | always | Calibration and standard-addition fits use the same notation.
m/q | Slope and intercept of the linear fit | y = m x + q | response / concentration and response | FIGURE_RGB.png, FIGURE_CIELAB_DELTAE.png, 06_FITTING | fit rows present
MeanBG_* | Linearized local background intensity for the channel | background surface evaluated at the well and gamma linearized | dimensionless | RAW | always | Derived from inter-well background reconstruction.
MeanBG_Red/MeanBG_Green/MeanBG_Blue | Linearized local inter-well background intensity for each RGB channel | 2D background surface evaluated at the well and gamma-linearized | dimensionless | 04_RAW | always
MeanW_* | Linearized median well intensity for the channel | median ROI channel intensity after gamma linearization | dimensionless | RAW | always | Computed from selected well ROI pixels.
MeanW_Red/MeanW_Green/MeanW_Blue | Linearized median well intensity for each RGB channel | median ROI channel intensity after gamma linearization | dimensionless | 04_RAW | always
Method/Family/ComparableGroup/RankMode | Method-comparison identifiers and comparable-score grouping | text labels defining descriptor, family and score comparability | text | METHOD_COMPARISON.png, 07_METHOD_COMPARISON | method comparison present | Scores should be compared directly only within the same ComparableGroup.
n_points | Number of finite data points used by the fit | count of finite x,y pairs | count | FITTING | always | All finite points are retained by robust IRLS.
NClipPoints/ClipX/ClipDelta | Calibration clipping/baseline diagnostic fields | NClipPoints = number of adjusted points; ClipX = concentrations; ClipDelta = applied response adjustments | count, concentration, response | 06_FITTING | calibration present | Diagnostic only; clipping is not part of the common method score.
NReplicates | Number of replicate wells in a summarized group | count of wells in group | count | 05_REPLICATES_MEAN | always
NWellWarnings/NWellCritical | Counts of replicate wells with warning or critical optical QC | count over replicate wells | count | REPLICATES_MEAN | always
PAbs | Image-derived RGB pseudo-absorbance | PAbs = log10(I_BG / I_well) = -log10(I_well / I_BG) | dimensionless | FIGURE_RGB.png, RESULTS_CAPTION.txt, 04_RAW, 05_REPLICATES_MEAN, 06_FITTING, 07_METHOD_COMPARISON | always | Pseudo-absorbance is not assumed to be spectrophotometric absorbance.
PAbs_Blue | RGB pseudo-absorbance (blue channel) | log10(BG_blue/W_blue) = -log10(W_blue/BG_blue) | dimensionless | RAW, REPLICATES_MEAN, FITTING | always | Image-derived pseudo-absorbance with fixed full-background normalization.
PAbs_Green | RGB pseudo-absorbance (green channel) | log10(BG_green/W_green) = -log10(W_green/BG_green) | dimensionless | RAW, REPLICATES_MEAN, FITTING | always | Image-derived pseudo-absorbance with fixed full-background normalization.
PAbs_Red | RGB pseudo-absorbance (red channel) | log10(BG_red/W_red) = -log10(W_red/BG_red) | dimensionless | RAW, REPLICATES_MEAN, FITTING | always | Image-derived pseudo-absorbance with fixed full-background normalization.
PAbs_Red/PAbs_Green/PAbs_Blue | RGB pseudo-absorbance for the red, green and blue channels | log10(MeanBG_channel / MeanW_channel) | dimensionless | FIGURE_RGB.png, 04_RAW, 05_REPLICATES_MEAN, 06_FITTING, 07_METHOD_COMPARISON | always | Exported in standard RGB order even though OpenCV stores images internally as BGR.
path_length | Estimated optical path length | l_cm = (V_uL / A_bottom_mm2) / 10 | cm | OVERVIEW, FITTING | epsilon mode configured | Assumes 1 uL = 1 mm^3 and a flat-bottom well with nominal bottom area.
path_length_mm | Estimated liquid height | l_mm = V_uL / A_bottom_mm2 | mm | OVERVIEW, FITTING | epsilon mode configured | Converted to cm before Beer-Lambert-like concentration calculation.
Purpose | Short description of the sheet role | free-text description | text | 01_CONTENTS | always
q | Intercept of linear fit | y = m x + q | signal | FITTING | always | For standard addition, C0 is derived from m and q.
QCCritical | Group-level critical optical quality flag | 1 if FracWellCritical >= 0.50 | 0/1 | REPLICATES_MEAN | always | Critical flag; points remain reported and used.
QCFlagged | Group-level optical quality warning | 1 if any replicate has ImageWarning, any critical replicate, or median UsedFraction < 0.70 | 0/1 | REPLICATES_MEAN | always | Warning flag; points remain reported and used.
QCFlagged/QCCritical | Replicate-group QC warning and critical flags | rule-based aggregation of well-level warnings, critical wells and UsedFraction | 0/1 | 05_REPLICATES_MEAN | always
R2 | Coefficient of determination | 1 - SSE/SST | dimensionless | FIGURE_RGB.png, FIGURE_CIELAB_DELTAE.png, 06_FITTING, 07_METHOD_COMPARISON | fit rows present | Closer to 1 indicates better fit.
R2_cal/R2_std_mean/m_cal/m_std_mean | Calibration and standard-addition fit descriptors used in method comparison | R2_cal from calibration fit; R2_std_mean = mean standard-addition R2; m_cal and m_std_mean are fit slopes | mixed | METHOD_COMPARISON.png, 07_METHOD_COMPARISON | method comparison present
Rank | Workbook-level rank of each method | 1 = highest Score within the selected ComparableGroup | integer | METHOD_COMPARISON | always | Rows outside the selected ComparableGroup are not assigned a direct rank.
RankMode | Data basis available for ranking | calibration_plus_stdadd, calibration_only, stdadd_only, or unavailable | text | METHOD_COMPARISON | always
recovery_pct_* | Recovery relative to external reference | 100 x estimate / reference | % | FIGURE_RGB.png, METHOD_COMPARISON.png, 03_OVERVIEW, 07_METHOD_COMPARISON | reference values configured
rel_error_* | Relative error versus external reference | 100 x (estimate - reference) / reference | % | 03_OVERVIEW, 07_METHOD_COMPARISON | reference values configured
Reliability criteria | Rule used to flag a method as low reliability in method comparison | low if slope agreement < 0.5 OR R² < 0.8 OR ¦Δ¦ > threshold | rule | METHOD_COMPARISON | stdadd present | The Δ threshold is max(3×reference SD, 50% of the reference value), evaluated against the nearest available reference.
Reliability score | Overall report-level reliability score | starts at 50 and adds/subtracts rule-based penalties/bonuses for calibration availability, stored calibration, plate QC, critical wells, used fraction, empty-well QC, and ranking separation | 0-100 | OVERVIEW | always | Not a statistical probability; it is a heuristic audit score.
RGB linearization | Conversion of gamma-domain RGB values to an intensity-like scale | I_lin = (I/255)^gamma, gamma approximately 2.2 | dimensionless | PAbs calculation in 04_RAW and FIGURE_RGB.png | always | No automatic image correction is applied.
RGB to CIELAB | Diagnostic conversion from image RGB/BGR values to CIE L*a*b* descriptors | OpenCV BGR->Lab for 8-bit sRGB-like images; reference form: linearized sRGB -> XYZ using the standard sRGB-to-XYZ matrix and D65 reference white -> CIE L*a*b* | dimensionless | FIGURE_CIELAB_DELTAE.png, RAW_DATA_DETAILS_CAPTION.txt, 04_RAW, 05_REPLICATES_MEAN, DIAGNOSTICS | CIELAB/DeltaE outputs present | References: OpenCV color conversions; IEC 61966-2-1:1999 sRGB; CIE 1976 L*a*b*; CIE standard illuminant D65.
RMSE | Root mean squared error of the fit | sqrt(mean(residual^2)) | response unit | 06_FITTING | fit rows present | Interpret in the units of the fitted response.
Robust SD | Robust dispersion across replicates | 1.4826 × MAD, where MAD = median(¦x − median(x)¦) | same as variable | REPLICATES_MEAN | always | Computed within each replicate group.
Row | Plate row label | A-based alphabetical row label | well-row label | RAW, REPLICATES_MEAN | always
S0_calibration | Channel-wide low-signal offset estimated from raw calibration | max(0, −min(raw calibration response)) | signal | FITTING | RGB calibration | Estimated before forced-zero calibration and transferred to StdAdd/unknown data.
Score | Common method-ranking score | for calibration plus standard addition: SlopeAgreement^2 x sqrt(R2_cal x R2_std) x (1/LOQ) | 1/mM | FIGURE_RGB.png, METHOD_COMPARISON.png, 07_METHOD_COMPARISON | method comparison present | Expected/reference values, recovery, SNR and clipping are not used in this score.
ScoreFormula | Formula used to compute Score | text descriptor | text | METHOD_COMPARISON | always | Documents which score formula was applied to the row.
Selected | Workbook-level selected method flag | 1 for the highest-ranked row within the most informative ComparableGroup, otherwise 0 | 0/1 | METHOD_COMPARISON | always | Selection in the workbook is derived from method-comparison rank and not from expected/recovery.
Selected/Rank | Selected method flag and rank position | Selected = 1 for the selected method; Rank = score order within the highest common-factor group | 0/1 and integer | 07_METHOD_COMPARISON | method comparison present
Sheet | Workbook sheet name | worksheet name | text | 01_CONTENTS | always
sigma_cal | Noise estimate for calibration | SD at zero concentration or median calibration SD | signal | FITTING | calibration only or calibration present | Source is reported in sigma_source.
sigma_cal/sigma_source/SNR | Calibration noise estimate, its source and slope-to-noise ratio | SNR = ¦m¦ / sigma_cal | mixed | 06_FITTING, 07_METHOD_COMPARISON | calibration present | SNR is diagnostic only and is not used in the common score.
sigma_source | Source of sigma_cal | median_calibration_sd, blank_zero_calibration_sd, or unavailable | text | FITTING | calibration present
SignalT_* | Transmittance-like intensity ratio before logarithm | MeanW_* / MeanBG_* | dimensionless | RAW | always | PAbs_* = -log10(SignalT_*).
SignalT_Red/SignalT_Green/SignalT_Blue | Transmittance-like ratio before logarithmic conversion | SignalT_channel = MeanW_channel / MeanBG_channel | dimensionless | 04_RAW | always | PAbs_channel = -log10(SignalT_channel).
SlopeAgreement | Agreement between calibration and standard-addition slope magnitudes | min(¦m_cal¦, ¦m_std¦) / max(¦m_cal¦, ¦m_std¦) | dimensionless | METHOD_COMPARISON.png, 07_METHOD_COMPARISON | calibration plus standard addition present | Equals 1 when slopes are identical in magnitude; lower values indicate poorer agreement between calibration and standard addition.
SNR | Slope-to-noise ratio | ¦m¦ / sigma_cal | dimensionless | FITTING | calibration only or calibration present | Used only for calibration rows.
Type | Well role in the analytical workflow | C = calibration; A = standard addition; U = unknown; other labels may be reported when present | text | RAW, REPLICATES_MEAN | always
UsedFraction | Fraction of ROI core pixels retained after filtering | n_used / n_core | dimensionless | RAW, SUMMARY | always | Lower values indicate stronger filtering. Closer to 1 means less aggressive trimming.
UsedFraction_median | Median UsedFraction across replicate wells in a group | median(UsedFraction) | dimensionless | REPLICATES_MEAN | always
Value | Value associated with a metadata or overview item | reported value for the corresponding Field | mixed | 02_METADATA, 03_OVERVIEW | always
Well | Human-readable well identifier | Well = Row + Col | well label | RAW, REPLICATES_MEAN | always
well_bottom_area_mm2 | Nominal flat-bottom well area used for path-length calculation | pi × (d_bottom/2)^2 | mm^2 | OVERVIEW | epsilon mode configured | Derived from the selected plate geometry.