## 01_CONTENTS
Sheet | Purpose
01_CONTENTS | Index of diagnostic sheets.
02_BG_SAMPLES | Accepted inter-well background samples used to fit the BG surface.
03_BG_WELL_FIT | Predicted local background at each well.
04_WELL_ROBUST_STATS | Well-level robust pixel statistics and optical QC.
05_GEOMETRY_QC | Floor/mouth geometry quality-control descriptors.
06_WELL_BOTTOM | Detailed well-bottom and mouth geometry measurements.
07_PLATE_GEOMETRY | Nominal plate geometry parameters used by the analyzer.
08_EMPTY_WELLS | Empty-well diagnostic values when empty wells are present.
09_SPATIAL_DIAGNOSTICS | Spatial trends across row/column positions.
10_METHOD_COMPARISON | Cross-method diagnostic comparison using common score factors.
11_CIELAB_FITTING | CIELAB/DeltaE diagnostic fit rows.
12_LEGENDS | Definitions for diagnostic workbook fields and figures.

## 02_BG_SAMPLES
BG_Cell_Row | BG_Cell_Col | Associated_Wells | x | y | area | Red_median_raw | Green_median_raw | Blue_median_raw
0 | 0 | A1-A2-B1-B2 | 244.6628 | 179.5164 | 519 | 208 | 208 | 209
0 | 1 | A2-A3-B2-B3 | 341.2374 | 182.0129 | 695 | 206 | 203 | 210
0 | 2 | A3-A4-B3-B4 | 438.8307 | 183.2058 | 821 | 203 | 200 | 209
0 | 3 | A4-A5-B4-B5 | 536.5467 | 183.0772 | 803 | 201 | 198 | 207
0 | 4 | A5-A6-B5-B6 | 633.0509 | 182.9008 | 786 | 201 | 197 | 209
0 | 5 | A6-A7-B6-B7 | 728.5355 | 182.2934 | 760 | 200 | 195 | 205
0 | 6 | A7-A8-B7-B8 | 822.3625 | 182.8829 | 811 | 199 | 194 | 206
0 | 7 | A8-A9-B8-B9 | 917.3784 | 182.8705 | 880 | 198 | 194 | 206
0 | 8 | A9-A10-B9-B10 | 1011.5663 | 183.5028 | 897 | 196 | 193 | 204
0 | 9 | A10-A11-B10-B11 | 1105.8279 | 183.9717 | 883 | 197 | 191 | 204
0 | 10 | A11-A12-B11-B12 | 1200.6247 | 184.5036 | 834 | 195 | 190 | 200
1 | 0 | B1-B2-C1-C2 | 245.2342 | 278.2613 | 444 | 208 | 208 | 210
1 | 1 | B2-B3-C2-C3 | 342.7739 | 279.6954 | 522 | 205 | 202 | 208
1 | 2 | B3-B4-C3-C4 | 439.6993 | 280.206 | 602 | 204 | 198 | 208
1 | 3 | B4-B5-C4-C5 | 537.5279 | 279.3784 | 555 | 203 | 197 | 208
1 | 4 | B5-B6-C5-C6 | 631.458 | 280.1664 | 703 | 202 | 196 | 208
1 | 5 | B6-B7-C6-C7 | 727.7798 | 279.5342 | 672 | 199 | 194 | 206
1 | 6 | B7-B8-C7-C8 | 821.715 | 280.4431 | 765 | 200 | 194 | 206
1 | 7 | B8-B9-C8-C9 | 916.6023 | 280.3611 | 767 | 199 | 193 | 205
1 | 8 | B9-B10-C9-C10 | 1009.9323 | 279.7898 | 709 | 199 | 192 | 205
1 | 9 | B10-B11-C10-C11 | 1104.8017 | 280.7408 | 706 | 198 | 191 | 204
1 | 10 | B11-B12-C11-C12 | 1199.0775 | 280.8434 | 645 | 195 | 189 | 201
2 | 0 | C1-C2-D1-D2 | 217.1378 | 400.5638 | 392 | 205 | 206 | 204
2 | 1 | C2-C3-D2-D3 | 346.115 | 378.8375 | 400 | 205 | 202 | 209
2 | 2 | C3-C4-D3-D4 | 439.9881 | 378.3307 | 505 | 202 | 198 | 208
2 | 3 | C4-C5-D4-D5 | 536.3175 | 378.4709 | 567 | 203 | 197 | 207
2 | 4 | C5-C6-D5-D6 | 630.0131 | 376.8956 | 613 | 201 | 195 | 208
2 | 5 | C6-C7-D6-D7 | 726.6062 | 376.2196 | 551 | 200 | 194 | 206
2 | 6 | C7-C8-D7-D8 | 820.4185 | 376.6 | 650 | 199 | 193 | 205
2 | 7 | C8-C9-D8-D9 | 915.5016 | 376.4283 | 628 | 198 | 192 | 205
2 | 8 | C9-C10-D9-D10 | 1009.1543 | 377.1383 | 622 | 197 | 191 | 204
2 | 9 | C10-C11-D10-D11 | 1103.8342 | 377.5233 | 579 | 197 | 190 | 203
2 | 10 | C11-C12-D11-D12 | 1198.0478 | 378.1575 | 565 | 195 | 189 | 201
3 | 0 | D1-D2-E1-E2 | 216.2335 | 498.9467 | 394 | 205 | 206 | 203
3 | 1 | D2-D3-E2-E3 | 345.7279 | 475.7245 | 294 | 203 | 199 | 206
3 | 2 | D3-D4-E3-E4 | 442.0804 | 475.6726 | 336 | 201 | 195 | 205
3 | 3 | D4-D5-E4-E5 | 537.6866 | 476.1852 | 351 | 201 | 194 | 205
3 | 4 | D5-D6-E5-E6 | 630.9542 | 474.8892 | 415 | 200 | 194 | 203
3 | 5 | D6-D7-E6-E7 | 726.0559 | 474.8559 | 465 | 199 | 193 | 204
3 | 6 | D7-D8-E7-E8 | 819.9277 | 474.7309 | 498 | 199 | 192 | 205
3 | 7 | D8-D9-E8-E9 | 914.3597 | 473.9336 | 467 | 197 | 191 | 202
3 | 8 | D9-D10-E9-E10 | 1007.8366 | 474.6882 | 465 | 197 | 190 | 202
3 | 9 | D10-D11-E10-E11 | 1102.1748 | 474.9957 | 469 | 195 | 188 | 199
3 | 10 | D11-D12-E11-E12 | 1196.3201 | 475.4834 | 453 | 194 | 188 | 198
4 | 0 | E1-E2-F1-F2 | 216.1751 | 595.6958 | 457 | 205 | 206 | 204
4 | 1 | E2-E3-F2-F3 | 345.8613 | 574.479 | 238 | 202 | 197 | 205
4 | 2 | E3-E4-F3-F4 | 440.7415 | 573.2959 | 294 | 200 | 195 | 202
4 | 3 | E4-E5-F4-F5 | 536.4532 | 573.5801 | 331 | 199 | 193 | 203
4 | 4 | E5-E6-F5-F6 | 631.7778 | 572.9938 | 324 | 199 | 193 | 203
4 | 5 | E6-E7-F6-F7 | 725.6462 | 571.152 | 342 | 198 | 192 | 202
4 | 6 | E7-E8-F7-F8 | 820.2248 | 571.853 | 347 | 198 | 192 | 203
4 | 7 | E8-E9-F8-F9 | 912.943 | 572.0333 | 421 | 196 | 190 | 199
4 | 8 | E9-E10-F9-F10 | 1006.9674 | 571.644 | 368 | 196 | 188 | 199
4 | 9 | E10-E11-F10-F11 | 1100.807 | 572.0643 | 342 | 194 | 187 | 198
4 | 10 | E11-E12-F11-F12 | 1194.9568 | 572.2099 | 324 | 194 | 188 | 196
5 | 0 | F1-F2-G1-G2 | 214.9673 | 693.2837 | 490 | 205 | 205 | 203
5 | 1 | F2-F3-G2-G3 | 344.9392 | 671.2928 | 181 | 202 | 197 | 204
5 | 2 | F3-F4-G3-G4 | 440.7443 | 670.7045 | 176 | 201 | 194 | 202
5 | 3 | F4-F5-G4-G5 | 535.6255 | 671.7041 | 267 | 199 | 194 | 201
5 | 4 | F5-F6-G5-G6 | 630.5917 | 670.0417 | 240 | 198 | 191 | 201
5 | 5 | F6-F7-G6-G7 | 725.5584 | 671.0693 | 274 | 198 | 193 | 200
5 | 6 | F7-F8-G7-G8 | 817.3731 | 668.9388 | 327 | 197 | 191 | 199
5 | 7 | F8-F9-G8-G9 | 912.911 | 670.2822 | 326 | 196 | 191 | 198
5 | 8 | F9-F10-G9-G10 | 1005.6151 | 668.5324 | 278 | 195 | 189 | 198
5 | 9 | F10-F11-G10-G11 | 1097.9896 | 669.5751 | 193 | 194 | 188 | 195
5 | 10 | F11-F12-G11-G12 | 1193.5066 | 668.9559 | 227 | 194 | 188 | 196
6 | 0 | G1-G2-H1-H2 | 213.3598 | 790.6736 | 478 | 205 | 206 | 201
6 | 1 | G2-G3-H2-H3 | 344.3533 | 769.6133 | 150 | 203 | 198 | 205
6 | 2 | G3-G4-H3-H4 | 440.6393 | 768.1967 | 122 | 200 | 196 | 204
6 | 3 | G4-G5-H4-H5 | 534.6813 | 767.2527 | 182 | 199 | 194 | 201
6 | 4 | G5-G6-H5-H6 | 630.3406 | 764.7246 | 138 | 199 | 193 | 200
6 | 5 | G6-G7-H6-H7 | 724.8857 | 766.9029 | 175 | 197 | 191 | 199
6 | 6 | G7-G8-H7-H8 | 817.0813 | 767.177 | 209 | 197 | 191 | 199
6 | 7 | G8-G9-H8-H9 | 911.5935 | 766.8618 | 123 | 197 | 192 | 199
6 | 8 | G9-G10-H9-H10 | 1005.5571 | 768.2237 | 438 | 191.5 | 186 | 192
6 | 9 | G10-G11-H10-H11 | 1098.1195 | 764.5786 | 159 | 193 | 188 | 195
6 | 10 | G11-G12-H11-H12 | 1193.0533 | 766.5467 | 150 | 192.5 | 187 | 194

## 03_BG_WELL_FIT
Row | Col | Well | x | y | BG_Red_raw | BG_Green_raw | BG_Blue_raw
A | 1 | A1 | 185.8214 | 143.5893 | 208.0673 | 209.4029 | 209.9458
A | 2 | A2 | 283.4246 | 143.7599 | 206.5583 | 206.2485 | 209.8378
A | 3 | A3 | 381.0278 | 143.9306 | 205.1303 | 203.4141 | 209.5996
A | 4 | A4 | 478.631 | 144.1012 | 203.7831 | 200.8997 | 209.2312
A | 5 | A5 | 576.2341 | 144.2718 | 202.5168 | 198.7053 | 208.7324
A | 6 | A6 | 673.8373 | 144.4425 | 201.3314 | 196.8309 | 208.1035
A | 7 | A7 | 771.4405 | 144.6131 | 200.2268 | 195.2765 | 207.3442
A | 8 | A8 | 869.0436 | 144.7837 | 199.2032 | 194.0422 | 206.4548
A | 9 | A9 | 966.6469 | 144.9544 | 198.2604 | 193.1279 | 205.4351
A | 10 | A10 | 1064.25 | 145.125 | 197.3986 | 192.5336 | 204.2851
A | 11 | A11 | 1161.8531 | 145.2956 | 196.6176 | 192.2593 | 203.0049
A | 12 | A12 | 1259.4563 | 145.4663 | 195.9175 | 192.305 | 201.5944
B | 1 | B1 | 185.017 | 242.3991 | 207.4834 | 207.9113 | 209.1745
B | 2 | B2 | 282.6324 | 242.4966 | 205.9839 | 204.7742 | 209.0597
B | 3 | B3 | 380.2477 | 242.5941 | 204.5654 | 201.9572 | 208.8148
B | 4 | B4 | 477.8631 | 242.6916 | 203.2277 | 199.4602 | 208.4395
B | 5 | B5 | 575.4784 | 242.7891 | 201.9709 | 197.2833 | 207.934
B | 6 | B6 | 673.0938 | 242.8866 | 200.795 | 195.4264 | 207.2982
B | 7 | B7 | 770.7092 | 242.9841 | 199.7 | 193.8897 | 206.5321
B | 8 | B8 | 868.3245 | 243.0816 | 198.6859 | 192.673 | 205.6358
B | 9 | B9 | 965.9398 | 243.1791 | 197.7527 | 191.7763 | 204.6092
B | 10 | B10 | 1063.5552 | 243.2766 | 196.9004 | 191.1998 | 203.4523
B | 11 | B11 | 1161.1705 | 243.3742 | 196.129 | 190.9433 | 202.1651
B | 12 | B12 | 1258.7859 | 243.4716 | 195.4384 | 191.0069 | 200.7477
C | 1 | C1 | 184.2126 | 341.2089 | 206.8921 | 206.663 | 208.1765
C | 2 | C2 | 281.8401 | 341.2333 | 205.4022 | 203.5429 | 208.0554
C | 3 | C3 | 379.4677 | 341.2577 | 203.9932 | 200.7429 | 207.804
C | 4 | C4 | 477.0952 | 341.282 | 202.665 | 198.263 | 207.4223
C | 5 | C5 | 574.7228 | 341.3064 | 201.4178 | 196.1032 | 206.9103
C | 6 | C6 | 672.3503 | 341.3308 | 200.2514 | 194.2636 | 206.268
C | 7 | C7 | 769.9778 | 341.3552 | 199.166 | 192.744 | 205.4954
C | 8 | C8 | 867.6054 | 341.3795 | 198.1614 | 191.5446 | 204.5925
C | 9 | C9 | 965.233 | 341.4039 | 197.2378 | 190.6653 | 203.5594
C | 10 | C10 | 1062.8605 | 341.4283 | 196.395 | 190.1061 | 202.3959
C | 11 | C11 | 1160.488 | 341.4527 | 195.6332 | 189.8671 | 201.1021
C | 12 | C12 | 1258.1156 | 341.4771 | 194.9522 | 189.9481 | 199.6781
D | 1 | D1 | 183.4082 | 440.0187 | 206.2936 | 205.6581 | 206.9521
D | 2 | D2 | 281.0479 | 439.97 | 204.8132 | 202.5546 | 206.8249
D | 3 | D3 | 378.6877 | 439.9212 | 203.4137 | 199.7712 | 206.5673
D | 4 | D4 | 476.3274 | 439.8725 | 202.0951 | 197.308 | 206.1795
D | 5 | D5 | 573.9672 | 439.8237 | 200.8574 | 195.165 | 205.6614
D | 6 | D6 | 671.6069 | 439.775 | 199.7006 | 193.3422 | 205.0129
D | 7 | D7 | 769.2466 | 439.7262 | 198.6247 | 191.8396 | 204.2341
D | 8 | D8 | 866.8864 | 439.6774 | 197.6297 | 190.6571 | 203.3251
D | 9 | D9 | 964.5261 | 439.6287 | 196.7156 | 189.7948 | 202.2857
D | 10 | D10 | 1062.1658 | 439.58 | 195.8825 | 189.2526 | 201.116
D | 11 | D11 | 1159.8055 | 439.5312 | 195.1302 | 189.0306 | 199.8159
D | 12 | D12 | 1257.4453 | 439.4824 | 194.4588 | 189.1289 | 198.3856
E | 1 | E1 | 182.6037 | 538.8286 | 205.6877 | 204.8965 | 205.501
E | 2 | E2 | 280.2557 | 538.7067 | 204.2168 | 201.8092 | 205.3681
E | 3 | E3 | 377.9076 | 538.5848 | 202.8269 | 199.0422 | 205.1048
E | 4 | E4 | 475.5595 | 538.463 | 201.5178 | 196.5954 | 204.7112
E | 5 | E5 | 573.2114 | 538.3411 | 200.2897 | 194.4688 | 204.1872
E | 6 | E6 | 670.8634 | 538.2192 | 199.1425 | 192.6624 | 203.5329
E | 7 | E7 | 768.5153 | 538.0973 | 198.0762 | 191.1763 | 202.7483
E | 8 | E8 | 866.1672 | 537.9754 | 197.0908 | 190.0104 | 201.8334
E | 9 | E9 | 963.8192 | 537.8535 | 196.1863 | 189.1647 | 200.7881
E | 10 | E10 | 1061.4711 | 537.7316 | 195.3627 | 188.6392 | 199.6125
E | 11 | E11 | 1159.123 | 537.6097 | 194.62 | 188.434 | 198.3065
E | 12 | E12 | 1256.7749 | 537.4879 | 193.9582 | 188.549 | 196.8702
F | 1 | F1 | 181.7993 | 637.6384 | 205.0745 | 204.3782 | 203.8234
F | 2 | F2 | 279.4634 | 637.4434 | 203.6132 | 201.3068 | 203.6851
F | 3 | F3 | 377.1276 | 637.2483 | 202.2328 | 198.5558 | 203.4164
F | 4 | F4 | 474.7917 | 637.0533 | 200.9333 | 196.125 | 203.0173
F | 5 | F5 | 572.4558 | 636.8583 | 199.7148 | 194.0144 | 202.4879
F | 6 | F6 | 670.1199 | 636.6633 | 198.5771 | 192.2242 | 201.8281
F | 7 | F7 | 767.7841 | 636.4683 | 197.5204 | 190.7542 | 201.0379
F | 8 | F8 | 865.4481 | 636.2732 | 196.5446 | 189.6045 | 200.1174
F | 9 | F9 | 963.1123 | 636.0782 | 195.6497 | 188.7751 | 199.0666
F | 10 | F10 | 1060.7764 | 635.8832 | 194.8357 | 188.266 | 197.8854
F | 11 | F11 | 1158.4406 | 635.6882 | 194.1026 | 188.0771 | 196.5738
F | 12 | F12 | 1256.1046 | 635.4932 | 193.4505 | 188.2086 | 195.1319
G | 1 | G1 | 180.9949 | 736.4481 | 204.4539 | 204.1032 | 201.9193
G | 2 | G2 | 278.6712 | 736.18 | 203.0022 | 201.0474 | 201.7759
G | 3 | G3 | 376.3475 | 735.9119 | 201.6314 | 198.312 | 201.502
G | 4 | G4 | 474.0239 | 735.6437 | 200.3415 | 195.8968 | 201.0978
G | 5 | G5 | 571.7001 | 735.3756 | 199.1326 | 193.802 | 200.5633
G | 6 | G6 | 669.3765 | 735.1074 | 198.0045 | 192.0275 | 199.8983
G | 7 | G7 | 767.0528 | 734.8393 | 196.9574 | 190.5734 | 199.103
G | 8 | G8 | 864.7291 | 734.5712 | 195.9912 | 189.4395 | 198.1773
G | 9 | G9 | 962.4054 | 734.303 | 195.1059 | 188.626 | 197.1212
G | 10 | G10 | 1060.0818 | 734.0349 | 194.3015 | 188.1329 | 195.9348
G | 11 | G11 | 1157.7581 | 733.7668 | 193.5781 | 187.9601 | 194.6179
G | 12 | G12 | 1255.4343 | 733.4986 | 192.9356 | 188.1075 | 193.1707
H | 1 | H1 | 180.1905 | 835.2579 | 203.8261 | 204.0716 | 199.7886
H | 2 | H2 | 277.879 | 834.9167 | 202.3839 | 201.031 | 199.6404
H | 3 | H3 | 375.5675 | 834.5754 | 201.0227 | 198.3108 | 199.3618
H | 4 | H4 | 473.256 | 834.2341 | 199.7424 | 195.9109 | 198.9528
H | 5 | H5 | 570.9445 | 833.8928 | 198.5431 | 193.8315 | 198.4135
H | 6 | H6 | 668.6329 | 833.5516 | 197.4246 | 192.0724 | 197.7437
H | 7 | H7 | 766.3215 | 833.2103 | 196.3871 | 190.6337 | 196.9435
H | 8 | H8 | 864.0099 | 832.869 | 195.4305 | 189.5154 | 196.0129
H | 9 | H9 | 961.6985 | 832.5278 | 194.5549 | 188.7175 | 194.952
H | 10 | H10 | 1059.387 | 832.1865 | 193.7601 | 188.2399 | 193.7606
H | 11 | H11 | 1157.0754 | 831.8453 | 193.0463 | 188.0828 | 192.4388
H | 12 | H12 | 1254.7639 | 831.504 | 192.4134 | 188.246 | 190.9867

## 04_WELL_ROBUST_STATS
Row | Col | Well | n_roi | n_core | n_used | used_fraction | highlight_fraction_roi | highlight_fraction_core | Gray_mean | Gray_median | Gray_sd | Gray_p10 | Gray_p25 | Gray_p50 | Gray_p75 | Gray_p90 | Gray_iqr | Purple_mean | Purple_median | Purple_sd | Purple_p10 | Purple_p25 | Purple_p50 | Purple_p75 | Purple_p90 | Purple_iqr | L_mean | L_median | L_sd | L_p10 | L_p25 | L_p50 | L_p75 | L_p90 | L_iqr | a_mean | a_median | a_sd | a_p10 | a_p25 | a_p50 | a_p75 | a_p90 | a_iqr | b_mean | b_median | b_sd | b_p10 | b_p25 | b_p50 | b_p75 | b_p90 | b_iqr | BrightExcludedFraction | BrightExcludedMeanGray | BrightExcessMeanGray | HighlightIndex | is_image_quality_warning | warning_reason | Red_mean | Green_mean | Blue_mean | Red_median | Green_median | Blue_median | Red_sd | Green_sd | Blue_sd | Red_p10 | Green_p10 | Blue_p10 | Red_p25 | Green_p25 | Blue_p25 | Red_p50 | Green_p50 | Blue_p50 | Red_p75 | Green_p75 | Blue_p75 | Red_p90 | Green_p90 | Blue_p90 | Red_iqr | Green_iqr | Blue_iqr
A | 1 | A1 | 2082 | 1442 | 1442 | 1 | 0 | 0 | 209.4723 | 210 | 0.71728 | 209 | 209 | 210 | 210 | 210 | 1 | -2.862 | -2.5 | 0.48074 | -3.5 | -3.5 | -2.5 | -2.5 | -2.5 | 1 | 84.107 | 84.3137 | 0.28016 | 83.9216 | 83.9216 | 84.3137 | 84.3137 | 84.3137 | 0.39216 | -1.362 | -1 | 0.48074 | -2 | -2 | -1 | -1 | -1 | 1 | 1.362 | 1 | 0.48074 | 1 | 1 | 1 | 2 | 2 | 1 | 0.036755 | 211 | 1 | 0.036755 | 0 |  | 208.4723 | 210.4723 | 206.7483 | 209 | 211 | 207 | 0.71728 | 0.71728 | 1.2185 | 208 | 210 | 205 | 208 | 210 | 206 | 209 | 211 | 207 | 209 | 211 | 208 | 209 | 211 | 208 | 1 | 1 | 2
A | 2 | A2 | 2206 | 1554 | 1554 | 1 | 0 | 0 | 174.9421 | 175 | 0.91871 | 174 | 174 | 175 | 175 | 176 | 1 | 17.5241 | 18 | 1.3428 | 16 | 18 | 18 | 18 | 18.5 | 0 | 70.9055 | 70.9804 | 0.30965 | 70.5882 | 70.5882 | 70.9804 | 70.9804 | 71.3726 | 0.39216 | 8.0534 | 8 | 0.77158 | 7 | 8 | 8 | 8 | 9 | 0 | -12.3559 | -13 | 0.8621 | -13 | -13 | -13 | -12 | -11 | 1 | 0.025097 | 177.1026 | 1.1026 | 0.027671 | 0 |  | 178.0116 | 169.2317 | 195.5 | 178 | 169 | 196 | 1.0984 | 1.1463 | 1.2865 | 177 | 168 | 194 | 178 | 169 | 195 | 178 | 169 | 196 | 179 | 170 | 196 | 179 | 171 | 197 | 1 | 1 | 1
A | 3 | A3 | 2365 | 1691 | 1691 | 1 | 0 | 0 | 169.0426 | 169 | 0.72689 | 168 | 169 | 169 | 169 | 170 | 0 | 22.5071 | 22.5 | 1.1577 | 21.5 | 21.5 | 22.5 | 23.5 | 24.5 | 2 | 68.5959 | 68.6274 | 0.32364 | 68.2353 | 68.2353 | 68.6274 | 68.6274 | 69.0196 | 0.39215 | 10.8445 | 11 | 0.77616 | 10 | 10 | 11 | 11 | 12 | 1 | -15.5263 | -16 | 0.68784 | -16 | -16 | -16 | -15 | -15 | 1 | 0.024246 | 171.0244 | 1.0244 | 0.024837 | 0 |  | 173.5098 | 161.78 | 195.0645 | 173 | 162 | 195 | 1.0119 | 0.97549 | 1.1937 | 172 | 161 | 194 | 173 | 161 | 194 | 173 | 162 | 195 | 174 | 162 | 196 | 175 | 163 | 196 | 1 | 1 | 2
A | 4 | A4 | 2512 | 1821 | 1821 | 1 | 0 | 0 | 162.8402 | 163 | 0.90739 | 162 | 162 | 163 | 163 | 164 | 1 | 26.0406 | 26.5 | 1.2269 | 25 | 25 | 26.5 | 27.5 | 27.5 | 2.5 | 66.2037 | 66.2745 | 0.36371 | 65.8824 | 65.8824 | 66.2745 | 66.2745 | 66.6667 | 0.39216 | 12.5398 | 13 | 0.61767 | 12 | 12 | 13 | 13 | 13 | 1 | -17.3487 | -17 | 0.5535 | -18 | -18 | -17 | -17 | -17 | 1 | 0.048325 | 165.1023 | 1.1023 | 0.053267 | 0 |  | 169.129 | 154.3059 | 191.564 | 169 | 154 | 192 | 0.96096 | 1.2711 | 1.0792 | 168 | 153 | 190 | 169 | 153 | 191 | 169 | 154 | 192 | 170 | 155 | 192 | 170 | 156 | 193 | 1 | 2 | 1
A | 5 | A5 | 2645 | 1937 | 1937 | 1 | 0 | 0 | 159.0991 | 159 | 0.90367 | 158 | 159 | 159 | 160 | 160 | 1 | 29.985 | 30 | 1.083 | 29 | 30 | 30 | 30 | 32 | 0 | 64.6901 | 64.7059 | 0.34344 | 64.3137 | 64.3137 | 64.7059 | 65.098 | 65.098 | 0.78431 | 14.9055 | 15 | 0.65859 | 14 | 15 | 15 | 15 | 16 | 0 | -18.9742 | -19 | 0.82344 | -19 | -19 | -19 | -19 | -18 | 0 | 0.045431 | 161.2045 | 1.2045 | 0.054724 | 0 |  | 167.1064 | 149.0036 | 190.8709 | 167 | 149 | 191 | 0.99329 | 1.0095 | 1.3099 | 166 | 148 | 189 | 166 | 148 | 190 | 167 | 149 | 191 | 168 | 150 | 192 | 168 | 150 | 192 | 2 | 2 | 2
A | 6 | A6 | 2774 | 2053 | 2053 | 1 | 0 | 0 | 156.9795 | 157 | 0.65603 | 156 | 157 | 157 | 157 | 158 | 0 | 32.791 | 32.5 | 1.6175 | 30 | 32 | 32.5 | 34 | 34.5 | 2 | 63.5764 | 63.5294 | 0.30045 | 63.1373 | 63.5294 | 63.5294 | 63.5294 | 63.9216 | 0 | 16.3415 | 16 | 0.80678 | 15 | 16 | 16 | 17 | 17 | 1 | -20.7233 | -21 | 0.7669 | -21 | -21 | -21 | -21 | -19 | 0 | 0.019484 | 159.025 | 1.025 | 0.019971 | 0 |  | 165.9079 | 145.6522 | 190.9786 | 166 | 145 | 191 | 1.2038 | 1.0349 | 1.2419 | 165 | 145 | 189 | 165 | 145 | 191 | 166 | 145 | 191 | 167 | 146 | 192 | 167 | 147 | 192 | 2 | 1 | 1
A | 7 | A7 | 2872 | 2137 | 2137 | 1 | 0 | 0 | 152.8095 | 153 | 0.64458 | 152 | 152 | 153 | 153 | 153 | 1 | 36.1226 | 35.5 | 1.462 | 34.5 | 35.5 | 35.5 | 37 | 38 | 1.5 | 62.0738 | 61.9608 | 0.33754 | 61.5686 | 61.9608 | 61.9608 | 62.3529 | 62.3529 | 0.39216 | 18.2302 | 18 | 0.82866 | 17 | 18 | 18 | 19 | 19 | 1 | -22.5929 | -22 | 1.0226 | -24 | -24 | -22 | -22 | -21 | 2 | 0.080487 | 154.1686 | 1.1686 | 0.094057 | 0 |  | 163.0037 | 140.4296 | 190.1006 | 163 | 140 | 190 | 0.9953 | 0.98672 | 1.3473 | 162 | 139 | 188 | 163 | 140 | 190 | 163 | 140 | 190 | 164 | 141 | 191 | 164 | 141 | 192 | 1 | 1 | 1
A | 8 | A8 | 2949 | 2205 | 2205 | 1 | 0 | 0 | 149.3678 | 149 | 0.91883 | 148 | 149 | 149 | 150 | 150 | 1 | 37.9068 | 38 | 1.1681 | 37 | 37.5 | 38 | 38 | 40 | 0.5 | 60.7023 | 60.7843 | 0.32445 | 60.3922 | 60.3922 | 60.7843 | 60.7843 | 61.1765 | 0.39215 | 19.0399 | 19 | 0.57963 | 19 | 19 | 19 | 19 | 20 | 0 | -24.088 | -24 | 0.60433 | -25 | -24 | -24 | -24 | -24 | 0 | 0.084807 | 151.2674 | 1.2674 | 0.10748 | 0 |  | 159.9324 | 136.3456 | 188.5723 | 160 | 136 | 188 | 0.96324 | 1.0925 | 1.4126 | 159 | 135 | 187 | 159 | 136 | 188 | 160 | 136 | 188 | 160 | 137 | 189 | 161 | 138 | 191 | 1 | 1 | 1
A | 9 | A9 | 2980 | 2233 | 2233 | 1 | 0 | 0 | 146.5777 | 146 | 1.072 | 145 | 146 | 146 | 147 | 148 | 1 | 39.0275 | 40 | 1.4945 | 37.5 | 38 | 40 | 40 | 40 | 2 | 59.5218 | 59.6078 | 0.44255 | 59.2157 | 59.2157 | 59.6078 | 59.6078 | 60 | 0.39216 | 19.7976 | 20 | 0.85102 | 19 | 19 | 20 | 20 | 21 | 1 | -24.5446 | -25 | 0.75308 | -25 | -25 | -25 | -24 | -24 | 1 | 0.045678 | 149.2843 | 1.2843 | 0.058665 | 0 |  | 157.4577 | 133.0331 | 186.6637 | 157 | 133 | 187 | 1.095 | 1.3911 | 1.5902 | 156 | 132 | 185 | 157 | 132 | 186 | 157 | 133 | 187 | 158 | 134 | 187 | 159 | 135 | 188 | 1 | 2 | 1
A | 10 | A10 | 2950 | 2203 | 2203 | 1 | 0 | 0 | 144.7068 | 144 | 1.1406 | 144 | 144 | 144 | 145 | 146 | 1 | 39.352 | 40 | 1.5465 | 37.5 | 38.5 | 40 | 40 | 40 | 1.5 | 58.7628 | 58.8235 | 0.49088 | 58.4314 | 58.4314 | 58.8235 | 58.8235 | 59.2157 | 0.39216 | 19.8684 | 20 | 0.79331 | 19 | 20 | 20 | 20 | 21 | 0 | -24.7603 | -25 | 0.71619 | -25 | -25 | -25 | -25 | -24 | 0 | 0.060372 | 147.7368 | 1.7368 | 0.10486 | 0 |  | 155.5175 | 131.0231 | 185.2329 | 155 | 131 | 185 | 1.2729 | 1.4659 | 1.5824 | 154 | 130 | 183 | 155 | 130 | 185 | 155 | 131 | 185 | 156 | 132 | 186 | 157 | 133 | 187 | 1 | 2 | 1
A | 11 | A11 | 2871 | 2129 | 2129 | 1 | 0 | 0 | 142.5359 | 142 | 1.0069 | 142 | 142 | 142 | 143 | 144 | 1 | 39.3138 | 40 | 1.7246 | 37 | 38 | 40 | 40 | 41 | 2 | 57.9521 | 58.0392 | 0.45047 | 57.6471 | 57.6471 | 58.0392 | 58.0392 | 58.4314 | 0.39215 | 20.4965 | 21 | 0.97435 | 19 | 20 | 21 | 21 | 21 | 1 | -24.6238 | -25 | 0.96354 | -25 | -25 | -25 | -24 | -23 | 1 | 0.10991 | 144.5769 | 1.5769 | 0.17332 | 0 |  | 153.5806 | 128.8666 | 182.7802 | 153 | 129 | 183 | 1.2374 | 1.3874 | 1.7471 | 153 | 128 | 181 | 153 | 128 | 182 | 153 | 129 | 183 | 154 | 129 | 184 | 155 | 130 | 184 | 1 | 1 | 2
A | 12 | A12 | 2745 | 2007 | 2007 | 1 | 0 | 0 | 141.7185 | 141 | 1.462 | 141 | 141 | 141 | 142 | 143 | 1 | 35.3739 | 35.5 | 2.4459 | 31 | 34.5 | 35.5 | 37.5 | 37.5 | 3 | 57.7868 | 57.6471 | 0.63925 | 57.2549 | 57.2549 | 57.6471 | 58.0392 | 58.4314 | 0.78431 | 17.9322 | 18 | 1.3561 | 16 | 18 | 18 | 19 | 19 | 1 | -23.0164 | -23 | 1.2698 | -24 | -24 | -23 | -23 | -21 | 1 | 0.064275 | 146.0388 | 3.0388 | 0.19532 | 0 |  | 151.1136 | 129.6607 | 178.9557 | 151 | 129 | 179 | 1.1315 | 2.2026 | 1.8302 | 150 | 128 | 177 | 151 | 128 | 178 | 151 | 129 | 179 | 152 | 130 | 180 | 152 | 132 | 181 | 1 | 2 | 2
B | 1 | B1 | 1996 | 1371 | 1371 | 1 | 0 | 0 | 209.9942 | 210 | 0.74381 | 209 | 210 | 210 | 210 | 211 | 0 | -2.7035 | -2.5 | 0.61799 | -3.5 | -3 | -2.5 | -2.5 | -2.5 | 0.5 | 84.3134 | 84.3137 | 0.28685 | 83.9216 | 84.3137 | 84.3137 | 84.3137 | 84.7059 | 0 | -1.1364 | -1 | 0.47836 | -2 | -1 | -1 | -1 | -1 | 0 | 1.2684 | 1 | 0.4433 | 1 | 1 | 1 | 2 | 2 | 1 | 0.013858 | 212.2105 | 1.2105 | 0.016776 | 0 |  | 209.1036 | 210.9839 | 207.4573 | 209 | 211 | 208 | 0.87014 | 0.79121 | 1.1746 | 208 | 210 | 206 | 209 | 211 | 207 | 209 | 211 | 208 | 210 | 211 | 208 | 210 | 212 | 209 | 1 | 0 | 1
B | 2 | B2 | 2108 | 1469 | 1469 | 1 | 0 | 0 | 174.26 | 174 | 0.94993 | 173 | 174 | 174 | 175 | 175 | 1 | 18.6161 | 18.5 | 1.3782 | 17.5 | 18 | 18.5 | 20 | 20.5 | 2 | 70.5816 | 70.5882 | 0.42406 | 70.1961 | 70.1961 | 70.5882 | 70.9804 | 70.9804 | 0.78431 | 8.7202 | 9 | 0.90714 | 8 | 8 | 9 | 9 | 10 | 1 | -13.0361 | -13 | 1.0437 | -14 | -14 | -13 | -13 | -11 | 1 | 0.079646 | 176.1538 | 1.1538 | 0.091899 | 0 |  | 177.6576 | 168.1157 | 195.806 | 178 | 168 | 196 | 0.91117 | 1.2824 | 1.3496 | 177 | 167 | 194 | 177 | 167 | 195 | 178 | 168 | 196 | 178 | 169 | 197 | 179 | 170 | 197 | 1 | 2 | 2
B | 3 | B3 | 2233 | 1580 | 1580 | 1 | 0 | 0 | 168.0987 | 168 | 0.87416 | 167 | 168 | 168 | 169 | 169 | 1 | 23.9165 | 24.5 | 0.99666 | 22.5 | 23.5 | 24.5 | 24.5 | 24.5 | 1 | 68.2638 | 68.2353 | 0.33301 | 67.8431 | 68.2353 | 68.2353 | 68.2353 | 68.6274 | 0 | 11.6652 | 12 | 0.57491 | 11 | 11 | 12 | 12 | 12 | 1 | -15.9278 | -16 | 0.35391 | -16 | -16 | -16 | -16 | -16 | 0 | 0.053165 | 170.3214 | 1.3214 | 0.070253 | 0 |  | 173.5797 | 160.3411 | 194.9354 | 174 | 160 | 195 | 0.9678 | 1.1013 | 1.0282 | 172 | 159 | 194 | 173 | 160 | 194 | 174 | 160 | 195 | 174 | 161 | 195 | 175 | 162 | 196 | 1 | 1 | 1
B | 4 | B4 | 2362 | 1691 | 1691 | 1 | 0 | 0 | 162.1413 | 162 | 0.89019 | 161 | 162 | 162 | 163 | 163 | 1 | 27.7386 | 27.5 | 1.1095 | 26.5 | 27 | 27.5 | 29 | 29 | 2 | 65.7935 | 65.8824 | 0.42364 | 65.4902 | 65.4902 | 65.8824 | 66.2745 | 66.2745 | 0.78432 | 13.3944 | 13 | 0.57043 | 13 | 13 | 13 | 14 | 14 | 1 | -17.7836 | -18 | 0.47467 | -18 | -18 | -18 | -18 | -17 | 0 | 0.060319 | 164.2157 | 1.2157 | 0.073329 | 0 |  | 169.3732 | 152.8143 | 191.7327 | 170 | 153 | 192 | 1.0209 | 1.1766 | 0.89283 | 168 | 152 | 191 | 169 | 152 | 191 | 170 | 153 | 192 | 170 | 154 | 192 | 170 | 154 | 193 | 1 | 2 | 1
B | 5 | B5 | 2474 | 1787 | 1787 | 1 | 0 | 0 | 158.6318 | 159 | 0.70342 | 158 | 158 | 159 | 159 | 159 | 1 | 31.4312 | 32 | 1.9202 | 30 | 30 | 32 | 34 | 34 | 4 | 64.3512 | 64.3137 | 0.38087 | 63.9216 | 63.9216 | 64.3137 | 64.7059 | 64.7059 | 0.78432 | 15.6827 | 16 | 0.98961 | 15 | 15 | 16 | 17 | 17 | 2 | -20.0963 | -21 | 1.1499 | -21 | -21 | -21 | -19 | -19 | 2 | 0.059317 | 160.0943 | 1.0943 | 0.064913 | 0 |  | 167.0011 | 147.901 | 191.6631 | 167 | 148 | 192 | 1.0298 | 1.33 | 1.4263 | 166 | 146 | 190 | 166 | 147 | 191 | 167 | 148 | 192 | 168 | 149 | 193 | 168 | 149 | 193 | 2 | 2 | 2
B | 6 | B6 | 2580 | 1879 | 1879 | 1 | 0 | 0 | 155.62 | 156 | 0.81932 | 155 | 155 | 156 | 156 | 157 | 1 | 34.5072 | 34.5 | 1.1894 | 32.5 | 34.5 | 34.5 | 35.5 | 35.5 | 1 | 63.007 | 63.1373 | 0.31566 | 62.7451 | 62.7451 | 63.1373 | 63.1373 | 63.5294 | 0.39216 | 17.3837 | 18 | 0.80213 | 16 | 17 | 18 | 18 | 18 | 1 | -21.4689 | -21 | 0.67036 | -22 | -22 | -21 | -21 | -21 | 1 | 0.10538 | 157.1869 | 1.1869 | 0.12507 | 0 |  | 165.1969 | 143.7924 | 191.4023 | 165 | 144 | 191 | 1.0229 | 1.0291 | 1.2934 | 164 | 143 | 190 | 165 | 143 | 191 | 165 | 144 | 191 | 166 | 144 | 192 | 166 | 145 | 193 | 1 | 1 | 1
B | 7 | B7 | 2654 | 1943 | 1943 | 1 | 0 | 0 | 151.9089 | 152 | 0.75904 | 151 | 151 | 152 | 152 | 153 | 1 | 37.5103 | 38 | 1.0671 | 35.5 | 37.5 | 38 | 38 | 38.5 | 0.5 | 61.5753 | 61.5686 | 0.33326 | 61.1765 | 61.5686 | 61.5686 | 61.5686 | 61.9608 | 0 | 18.8163 | 19 | 0.51716 | 18 | 19 | 19 | 19 | 19 | 0 | -23.9218 | -24 | 0.82946 | -25 | -24 | -24 | -24 | -23 | 0 | 0.035512 | 154.2609 | 1.2609 | 0.044776 | 0 |  | 162.2126 | 139.0659 | 190.9398 | 162 | 139 | 191 | 0.99619 | 0.93275 | 1.2788 | 161 | 138 | 189 | 162 | 139 | 190 | 162 | 139 | 191 | 163 | 139 | 191 | 163 | 140 | 193 | 1 | 0 | 1
B | 8 | B8 | 2712 | 1997 | 1997 | 1 | 0 | 0 | 148.999 | 149 | 0.88349 | 148 | 148 | 149 | 150 | 150 | 2 | 39.2949 | 40 | 1.3682 | 38 | 38 | 40 | 40 | 41 | 2 | 60.561 | 60.3922 | 0.32461 | 60 | 60.3922 | 60.3922 | 60.7843 | 60.7843 | 0.39215 | 19.7882 | 20 | 0.82466 | 19 | 19 | 20 | 20 | 21 | 1 | -24.6465 | -25 | 0.75255 | -25 | -25 | -25 | -24 | -24 | 1 | 0.044567 | 151.191 | 1.191 | 0.05308 | 0 |  | 160.2263 | 135.4667 | 189.297 | 160 | 135 | 189 | 1.1078 | 1.1081 | 1.5073 | 159 | 134 | 187 | 160 | 135 | 188 | 160 | 135 | 189 | 161 | 136 | 190 | 162 | 137 | 191 | 1 | 1 | 2
B | 9 | B9 | 2730 | 2010 | 2010 | 1 | 0 | 0 | 145.7244 | 146 | 0.86963 | 145 | 145 | 146 | 146 | 147 | 1 | 41.2813 | 41 | 1.6489 | 40 | 40 | 41 | 43 | 43 | 3 | 59.1843 | 59.2157 | 0.34954 | 58.8235 | 58.8235 | 59.2157 | 59.2157 | 59.6078 | 0.39215 | 21.0687 | 21 | 0.92273 | 20 | 21 | 21 | 22 | 22 | 1 | -25.4667 | -26 | 0.77936 | -26 | -26 | -26 | -25 | -25 | 1 | 0.028358 | 148.3158 | 1.3158 | 0.037313 | 0 |  | 157.8259 | 131.4428 | 187.6224 | 158 | 131 | 188 | 1.1971 | 1.1915 | 1.6261 | 156 | 130 | 186 | 157 | 131 | 186 | 158 | 131 | 188 | 159 | 132 | 189 | 159 | 133 | 189 | 2 | 1 | 3
B | 10 | B10 | 2710 | 1993 | 1993 | 1 | 0 | 0 | 143.1726 | 143 | 0.87543 | 142 | 143 | 143 | 144 | 144 | 1 | 42.1217 | 43 | 1.6412 | 40 | 41 | 43 | 43 | 43.5 | 2 | 58.14 | 58.0392 | 0.38264 | 57.6471 | 58.0392 | 58.0392 | 58.4314 | 58.4314 | 0.39216 | 21.5454 | 22 | 0.82874 | 20 | 21 | 22 | 22 | 22 | 1 | -26.2434 | -26 | 0.95295 | -27 | -27 | -26 | -26 | -25 | 1 | 0.060211 | 145.5 | 1.5 | 0.090316 | 0 |  | 155.5253 | 128.572 | 185.862 | 156 | 128 | 186 | 1.1365 | 1.2655 | 1.3609 | 154 | 127 | 184 | 155 | 128 | 185 | 156 | 128 | 186 | 156 | 129 | 186 | 157 | 130 | 187 | 1 | 1 | 1
B | 11 | B11 | 2647 | 1939 | 1939 | 1 | 0 | 0 | 140.7143 | 141 | 1.087 | 140 | 140 | 141 | 141 | 142 | 1 | 41.1821 | 41 | 1.5039 | 40 | 41 | 41 | 43 | 43 | 2 | 57.3773 | 57.2549 | 0.44435 | 56.8627 | 57.2549 | 57.2549 | 57.6471 | 58.0392 | 0.39216 | 21.1377 | 21 | 0.73052 | 21 | 21 | 21 | 22 | 22 | 1 | -25.491 | -25 | 1.0661 | -27 | -27 | -25 | -25 | -25 | 2 | 0.054152 | 143.8 | 1.8 | 0.097473 | 0 |  | 153.2393 | 126.5333 | 182.1913 | 153 | 126 | 182 | 1.0705 | 1.4176 | 1.5473 | 152 | 125 | 181 | 153 | 126 | 181 | 153 | 126 | 182 | 154 | 127 | 183 | 154 | 128 | 184 | 1 | 1 | 2
B | 12 | B12 | 2544 | 1846 | 1846 | 1 | 0 | 0 | 139.7226 | 140 | 1.5102 | 138 | 139 | 140 | 140 | 141 | 1 | 37.1774 | 38 | 1.7454 | 34.5 | 37 | 38 | 38 | 38 | 1 | 57.0983 | 56.8627 | 0.60262 | 56.4706 | 56.8627 | 56.8627 | 57.2549 | 57.6471 | 0.39215 | 19.2806 | 20 | 1.1005 | 18 | 19 | 20 | 20 | 20 | 1 | -23.6224 | -24 | 0.93988 | -24 | -24 | -24 | -23 | -22 | 1 | 0.069881 | 143.5891 | 2.5891 | 0.18093 | 0 |  | 150.2156 | 126.9404 | 178.0201 | 150 | 127 | 178 | 1.4019 | 1.9016 | 1.5992 | 149 | 125 | 176 | 149 | 126 | 177 | 150 | 127 | 178 | 151 | 127 | 179 | 152 | 129 | 180 | 2 | 1 | 2
C | 1 | C1 | 1886 | 1278 | 1278 | 1 | 0 | 0 | 209.9961 | 210 | 0.67567 | 209 | 210 | 210 | 210 | 211 | 0 | -2.804 | -2.5 | 0.45952 | -3.5 | -3.5 | -2.5 | -2.5 | -2.5 | 1 | 84.3122 | 84.3137 | 0.26497 | 83.9216 | 84.3137 | 84.3137 | 84.3137 | 84.7059 | 0 | -1.2692 | -1 | 0.44897 | -2 | -2 | -1 | -1 | -1 | 1 | 1.3435 | 1 | 0.51004 | 1 | 1 | 1 | 2 | 2 | 1 | 0.010172 | 212 | 1 | 0.010172 | 0 |  | 209.0532 | 210.9937 | 207.3263 | 209 | 211 | 208 | 0.68796 | 0.68084 | 1.2617 | 208 | 210 | 206 | 209 | 211 | 206 | 209 | 211 | 208 | 209 | 211 | 208 | 210 | 212 | 209 | 0 | 0 | 2
C | 2 | C2 | 1986 | 1365 | 1365 | 1 | 0 | 0 | 174.2645 | 174 | 0.93934 | 173 | 174 | 174 | 175 | 175 | 1 | 18.7696 | 18.5 | 1.1349 | 18 | 18 | 18.5 | 19 | 20.5 | 1 | 70.5931 | 70.5882 | 0.40332 | 70.1961 | 70.1961 | 70.5882 | 70.9804 | 70.9804 | 0.78431 | 8.8828 | 9 | 0.75971 | 8 | 8 | 9 | 9 | 10 | 1 | -12.8132 | -13 | 1.0172 | -14 | -14 | -13 | -12 | -12 | 2 | 0.090842 | 176.1855 | 1.1855 | 0.10769 | 0 |  | 177.9062 | 168.0725 | 195.778 | 178 | 168 | 196 | 0.79404 | 1.1802 | 1.4245 | 177 | 167 | 194 | 177 | 167 | 195 | 178 | 168 | 196 | 178 | 169 | 197 | 179 | 170 | 197 | 1 | 2 | 2
C | 3 | C3 | 2093 | 1460 | 1460 | 1 | 0 | 0 | 167.0274 | 167 | 0.93111 | 166 | 167 | 167 | 167 | 168 | 0 | 24.438 | 24.5 | 0.98503 | 23.5 | 23.5 | 24.5 | 24.5 | 26 | 1 | 67.7763 | 67.8431 | 0.41457 | 67.451 | 67.451 | 67.8431 | 67.8431 | 68.2353 | 0.39216 | 12.0856 | 12 | 0.55801 | 12 | 12 | 12 | 12 | 13 | 0 | -15.7486 | -16 | 0.62326 | -16 | -16 | -16 | -15 | -15 | 1 | 0.077397 | 169.3097 | 1.3097 | 0.10137 | 0 |  | 173.1671 | 158.8637 | 193.4363 | 173 | 159 | 194 | 1.0231 | 1.1321 | 1.4437 | 172 | 158 | 192 | 173 | 158 | 192 | 173 | 159 | 194 | 174 | 159 | 194 | 174 | 160 | 195 | 1 | 1 | 2
C | 4 | C4 | 2186 | 1539 | 1539 | 1 | 0 | 0 | 161.9922 | 162 | 0.92146 | 161 | 161 | 162 | 162 | 163 | 1 | 29.0455 | 29 | 0.91079 | 28 | 29 | 29 | 29 | 30 | 0 | 65.5712 | 65.4902 | 0.38395 | 65.098 | 65.4902 | 65.4902 | 65.8824 | 65.8824 | 0.39216 | 14.1774 | 14 | 0.59176 | 14 | 14 | 14 | 14 | 15 | 0 | -18.1267 | -18 | 0.70494 | -19 | -18 | -18 | -18 | -17 | 0 | 0.068226 | 164.2571 | 1.2571 | 0.08577 | 0 |  | 169.9227 | 152.0227 | 192.2138 | 170 | 152 | 192 | 1.086 | 0.98302 | 1.4773 | 169 | 151 | 191 | 169 | 151 | 191 | 170 | 152 | 192 | 170 | 153 | 193 | 171 | 153 | 194 | 1 | 2 | 2
C | 5 | C5 | 2289 | 1623 | 1623 | 1 | 0 | 0 | 158.5786 | 159 | 0.7571 | 158 | 158 | 159 | 159 | 159 | 1 | 33.3453 | 34 | 1.357 | 32 | 32 | 34 | 34.5 | 34.5 | 2.5 | 64.1736 | 64.3137 | 0.32335 | 63.9216 | 63.9216 | 64.3137 | 64.3137 | 64.7059 | 0.39216 | 16.5601 | 17 | 0.63201 | 16 | 16 | 17 | 17 | 17 | 1 | -20.9088 | -21 | 0.54087 | -21 | -21 | -21 | -21 | -21 | 0 | 0.069624 | 160.1062 | 1.1062 | 0.077018 | 0 |  | 167.7973 | 147.0179 | 192.9291 | 168 | 147 | 193 | 1.2114 | 1.06 | 0.87747 | 166 | 146 | 192 | 167 | 146 | 193 | 168 | 147 | 193 | 169 | 148 | 193 | 169 | 148 | 194 | 2 | 2 | 0
C | 6 | C6 | 2374 | 1700 | 1700 | 1 | 0 | 0 | 154.9735 | 155 | 0.83261 | 154 | 154 | 155 | 155 | 156 | 1 | 35.5329 | 35.5 | 1.3045 | 34.5 | 34.5 | 35.5 | 37 | 37 | 2.5 | 62.7606 | 62.7451 | 0.3193 | 62.3529 | 62.7451 | 62.7451 | 62.7451 | 63.1373 | 0 | 17.9176 | 18 | 0.91049 | 17 | 17 | 18 | 19 | 19 | 2 | -22.0488 | -22 | 1.0256 | -23 | -23 | -22 | -21 | -21 | 2 | 0.045882 | 157.1667 | 1.1667 | 0.053529 | 0 |  | 165.1318 | 142.6535 | 191.2412 | 165 | 143 | 191 | 0.95035 | 1.1351 | 1.4518 | 164 | 141 | 190 | 165 | 142 | 190 | 165 | 143 | 191 | 166 | 143 | 192 | 166 | 144 | 193 | 1 | 1 | 2
C | 7 | C7 | 2442 | 1760 | 1760 | 1 | 0 | 0 | 151.3028 | 151 | 0.7432 | 151 | 151 | 151 | 152 | 152 | 1 | 38.0991 | 38 | 1.1764 | 37 | 37.5 | 38 | 38 | 40 | 0.5 | 61.3264 | 61.1765 | 0.33527 | 61.1765 | 61.1765 | 61.1765 | 61.5686 | 61.5686 | 0.39216 | 19.1744 | 19 | 0.62424 | 19 | 19 | 19 | 19 | 20 | 0 | -24.0784 | -24 | 0.85543 | -25 | -24 | -24 | -24 | -23 | 0 | 0.059659 | 153.1429 | 1.1429 | 0.068182 | 0 |  | 162.1057 | 138.221 | 190.5347 | 162 | 138 | 190 | 0.85049 | 0.90448 | 1.6964 | 161 | 137 | 189 | 162 | 138 | 190 | 162 | 138 | 190 | 162 | 139 | 191 | 163 | 139 | 193 | 0 | 1 | 1
C | 8 | C8 | 2479 | 1791 | 1791 | 1 | 0 | 0 | 147.6209 | 148 | 0.8728 | 147 | 147 | 148 | 148 | 149 | 1 | 41.1253 | 41 | 1.3513 | 40 | 40 | 41 | 41.5 | 43 | 1.5 | 59.8678 | 60 | 0.36331 | 59.6078 | 59.6078 | 60 | 60 | 60.3922 | 0.39215 | 20.9687 | 21 | 0.77757 | 20 | 21 | 21 | 21 | 22 | 0 | -25.7711 | -26 | 0.74449 | -27 | -26 | -26 | -25 | -25 | 1 | 0.017867 | 150.2188 | 1.2188 | 0.021776 | 0 |  | 159.0826 | 133.4456 | 190.0592 | 159 | 134 | 190 | 1.0318 | 1.1474 | 1.4216 | 158 | 132 | 188 | 158 | 133 | 189 | 159 | 134 | 190 | 160 | 134 | 191 | 160 | 135 | 192 | 2 | 1 | 2
C | 9 | C9 | 2494 | 1804 | 1804 | 1 | 0 | 0 | 144.3686 | 144 | 1.0135 | 143 | 144 | 144 | 145 | 146 | 1 | 42.505 | 43 | 1.2672 | 41 | 41 | 43 | 43 | 43.5 | 2 | 58.6172 | 58.4314 | 0.42861 | 58.0392 | 58.4314 | 58.4314 | 58.8235 | 59.2157 | 0.39216 | 21.7073 | 22 | 0.67827 | 21 | 21 | 22 | 22 | 22 | 1 | -26.1086 | -26 | 0.75828 | -27 | -27 | -26 | -26 | -25 | 1 | 0.10809 | 146.3949 | 1.3949 | 0.15078 | 0 |  | 157.0915 | 129.6208 | 187.1602 | 157 | 129 | 187 | 1.1752 | 1.2883 | 1.3152 | 156 | 128 | 185 | 156 | 129 | 186 | 157 | 129 | 187 | 158 | 130 | 188 | 158 | 131 | 189 | 2 | 1 | 2
C | 10 | C10 | 2472 | 1786 | 1786 | 1 | 0 | 0 | 142.0974 | 142 | 1.0614 | 141 | 141 | 142 | 143 | 143 | 2 | 42.9258 | 43 | 1.0814 | 41 | 43 | 43 | 43.5 | 43.5 | 0.5 | 57.7441 | 57.6471 | 0.44829 | 57.2549 | 57.6471 | 57.6471 | 58.0392 | 58.4314 | 0.39215 | 21.939 | 22 | 0.56029 | 21 | 22 | 22 | 22 | 22 | 0 | -26.523 | -27 | 0.8598 | -27 | -27 | -27 | -26 | -25 | 1 | 0.087906 | 144.4204 | 1.4204 | 0.12486 | 0 |  | 154.9686 | 127.2268 | 185.3365 | 155 | 127 | 185 | 1.0733 | 1.2247 | 1.5253 | 154 | 126 | 184 | 154 | 127 | 185 | 155 | 127 | 185 | 156 | 128 | 186 | 156 | 129 | 187 | 2 | 1 | 1
C | 11 | C11 | 2424 | 1746 | 1746 | 1 | 0 | 0 | 140.134 | 140 | 0.99415 | 139 | 140 | 140 | 141 | 141 | 1 | 42.4376 | 43 | 1.4144 | 41 | 41 | 43 | 43.5 | 43.5 | 2.5 | 57.0117 | 56.8627 | 0.44641 | 56.4706 | 56.8627 | 56.8627 | 57.2549 | 57.6471 | 0.39215 | 21.7405 | 22 | 0.69111 | 21 | 21 | 22 | 22 | 22 | 1 | -26.5361 | -27 | 0.90088 | -27 | -27 | -27 | -26 | -25 | 1 | 0.088774 | 142.3935 | 1.3935 | 0.12371 | 0 |  | 152.7554 | 125.4588 | 183.0372 | 153 | 125 | 183 | 0.90353 | 1.3477 | 1.4276 | 152 | 124 | 181 | 152 | 125 | 182 | 153 | 125 | 183 | 153 | 126 | 184 | 154 | 127 | 184 | 1 | 1 | 2
C | 12 | C12 | 2348 | 1675 | 1675 | 1 | 0 | 0 | 138.8245 | 138 | 1.5213 | 137 | 138 | 138 | 140 | 141 | 2 | 37.5997 | 38 | 2.0925 | 35.5 | 37.5 | 38 | 38 | 40 | 0.5 | 56.6701 | 56.4706 | 0.66188 | 56.0784 | 56.0784 | 56.4706 | 56.8627 | 57.2549 | 0.78431 | 19.5821 | 20 | 1.2932 | 18 | 19 | 20 | 20 | 21 | 1 | -23.9749 | -24 | 1.2011 | -26 | -24 | -24 | -24 | -23 | 0 | 0.10925 | 142 | 2 | 0.21851 | 0 |  | 149.3893 | 125.8693 | 177.5487 | 149 | 125 | 177 | 1.2468 | 2.1201 | 1.6564 | 148 | 124 | 175 | 149 | 125 | 177 | 149 | 125 | 177 | 150 | 127 | 179 | 151 | 129 | 179 | 1 | 2 | 2
D | 1 | D1 | 1771 | 1177 | 1177 | 1 | 0 | 0 | 208.7247 | 209 | 0.80009 | 208 | 209 | 209 | 209 | 209 | 0 | -2.4163 | -2.5 | 0.95943 | -3.5 | -3 | -2.5 | -1 | -1 | 2 | 83.8506 | 83.9216 | 0.25633 | 83.5294 | 83.9216 | 83.9216 | 83.9216 | 83.9216 | 0 | -0.85811 | -1 | 0.6056 | -2 | -1 | -1 | 0 | 0 | 1 | 1.452 | 1 | 0.54674 | 1 | 1 | 1 | 2 | 2 | 1 | 0.074766 | 210.0227 | 1.0227 | 0.076466 | 0 |  | 208.2931 | 209.486 | 205.8462 | 208 | 210 | 207 | 0.90156 | 0.83809 | 1.5389 | 207 | 208 | 203 | 208 | 209 | 205 | 208 | 210 | 207 | 209 | 210 | 207 | 209 | 210 | 207 | 1 | 1 | 2
D | 2 | D2 | 1854 | 1251 | 1251 | 1 | 0 | 0 | 173.446 | 173 | 0.75742 | 173 | 173 | 173 | 174 | 174 | 1 | 19.6675 | 20.5 | 1.3802 | 17.5 | 18.5 | 20.5 | 20.5 | 21.5 | 2 | 70.2412 | 70.1961 | 0.36273 | 69.8039 | 69.8039 | 70.1961 | 70.5882 | 70.5882 | 0.78431 | 9.4452 | 10 | 0.73157 | 8 | 9 | 10 | 10 | 10 | 1 | -13.1343 | -14 | 1.3343 | -15 | -14 | -14 | -12 | -11 | 2 | 0.073541 | 175.0217 | 1.0217 | 0.07514 | 0 |  | 177.55 | 166.8953 | 195.5755 | 177 | 167 | 196 | 0.83121 | 1.0085 | 1.8176 | 177 | 166 | 193 | 177 | 166 | 194 | 177 | 167 | 196 | 178 | 168 | 197 | 179 | 168 | 198 | 1 | 2 | 3
D | 3 | D3 | 1945 | 1330 | 1330 | 1 | 0 | 0 | 166.4767 | 166 | 0.73899 | 166 | 166 | 166 | 167 | 167 | 1 | 25.0951 | 25 | 1.039 | 24 | 24.5 | 25 | 26 | 26 | 1.5 | 67.4693 | 67.451 | 0.39541 | 67.0588 | 67.0588 | 67.451 | 67.8431 | 67.8431 | 0.78432 | 12.3872 | 12 | 0.6512 | 12 | 12 | 12 | 13 | 13 | 1 | -15.9857 | -16 | 0.45631 | -16 | -16 | -16 | -16 | -15 | 0 | 0.074436 | 168.101 | 1.101 | 0.081955 | 0 |  | 172.9135 | 158.0391 | 193.3549 | 173 | 158 | 193 | 0.93709 | 1.1013 | 0.97156 | 172 | 157 | 192 | 173 | 157 | 193 | 173 | 158 | 193 | 173 | 159 | 194 | 174 | 159 | 195 | 0 | 2 | 1
D | 4 | D4 | 2031 | 1405 | 1405 | 1 | 0 | 0 | 161.148 | 161 | 0.59986 | 161 | 161 | 161 | 161 | 162 | 0 | 28.9907 | 29.5 | 1.2636 | 27.5 | 27.5 | 29.5 | 30 | 30 | 2.5 | 65.424 | 65.4902 | 0.28413 | 65.098 | 65.098 | 65.4902 | 65.4902 | 65.8824 | 0.39216 | 14.4064 | 15 | 0.74005 | 13 | 14 | 15 | 15 | 15 | 1 | -18.6085 | -19 | 0.72731 | -19 | -19 | -19 | -18 | -18 | 1 | 0.015658 | 163.2727 | 1.2727 | 0.019929 | 0 |  | 168.5559 | 151.4498 | 192.3253 | 169 | 151 | 193 | 0.7562 | 0.77704 | 1.5285 | 168 | 151 | 191 | 168 | 151 | 191 | 169 | 151 | 193 | 169 | 152 | 193 | 169 | 152 | 194 | 1 | 1 | 2
D | 5 | D5 | 2115 | 1474 | 1474 | 1 | 0 | 0 | 157.386 | 157 | 0.80618 | 157 | 157 | 157 | 158 | 158 | 1 | 32.8721 | 32.5 | 1.2654 | 32 | 32 | 32.5 | 34 | 34.5 | 2 | 63.6986 | 63.5294 | 0.32208 | 63.5294 | 63.5294 | 63.5294 | 63.9216 | 63.9216 | 0.39216 | 16.2985 | 16 | 0.67677 | 16 | 16 | 16 | 17 | 17 | 1 | -21.1377 | -21 | 0.47692 | -22 | -21 | -21 | -21 | -21 | 0 | 0.065807 | 159.2784 | 1.2784 | 0.084125 | 0 |  | 165.8508 | 146.1052 | 192.1038 | 166 | 146 | 192 | 1.2296 | 1.0865 | 1.0478 | 165 | 145 | 191 | 165 | 146 | 192 | 166 | 146 | 192 | 167 | 147 | 193 | 168 | 147 | 193 | 2 | 1 | 1
D | 6 | D6 | 2177 | 1528 | 1528 | 1 | 0 | 0 | 154.7997 | 155 | 0.71595 | 154 | 154 | 155 | 155 | 156 | 1 | 34.946 | 35.5 | 1.1141 | 34.5 | 34.5 | 35.5 | 35.5 | 35.5 | 1 | 62.7215 | 62.7451 | 0.25713 | 62.3529 | 62.7451 | 62.7451 | 62.7451 | 63.1373 | 0 | 17.5255 | 18 | 0.73236 | 17 | 17 | 18 | 18 | 18 | 1 | -21.5733 | -22 | 0.74775 | -22 | -22 | -22 | -21 | -21 | 1 | 0.1034 | 156.1076 | 1.1076 | 0.11453 | 0 |  | 164.7264 | 142.7945 | 190.7546 | 165 | 143 | 191 | 0.89518 | 0.83375 | 1.4573 | 164 | 142 | 189 | 164 | 142 | 190 | 165 | 143 | 191 | 165 | 143 | 192 | 166 | 144 | 192 | 1 | 1 | 2
D | 7 | D7 | 2234 | 1578 | 1578 | 1 | 0 | 0 | 150.0488 | 150 | 0.7701 | 149 | 150 | 150 | 150 | 151 | 0 | 38.1461 | 38 | 1.3678 | 36.5 | 37.5 | 38 | 38.5 | 40 | 1 | 60.8626 | 60.7843 | 0.32515 | 60.3922 | 60.7843 | 60.7843 | 60.7843 | 61.1765 | 0 | 19.1502 | 19 | 0.67916 | 18 | 19 | 19 | 19 | 20 | 0 | -24.2548 | -24 | 0.90078 | -25 | -25 | -24 | -24 | -23 | 1 | 0.038657 | 152.2623 | 1.2623 | 0.048796 | 0 |  | 160.5228 | 136.9335 | 189.6362 | 161 | 137 | 189 | 0.93535 | 1.093 | 1.6299 | 159 | 136 | 188 | 160 | 136 | 189 | 161 | 137 | 189 | 161 | 137 | 191 | 161 | 138 | 192 | 1 | 1 | 2
D | 8 | D8 | 2275 | 1613 | 1613 | 1 | 0 | 0 | 146.7061 | 147 | 0.9228 | 146 | 146 | 147 | 147 | 148 | 1 | 41.4761 | 41 | 1.4259 | 40 | 41 | 41 | 43 | 43 | 2 | 59.5643 | 59.6078 | 0.36483 | 59.2157 | 59.2157 | 59.6078 | 59.6078 | 60 | 0.39216 | 21.2697 | 21 | 0.70697 | 21 | 21 | 21 | 22 | 22 | 1 | -25.3881 | -25 | 0.78251 | -26 | -26 | -25 | -25 | -24 | 1 | 0.032858 | 149.434 | 1.434 | 0.047117 | 0 |  | 159.2455 | 132.3726 | 188.452 | 159 | 132 | 189 | 1.1063 | 1.2082 | 1.6052 | 158 | 131 | 186 | 159 | 132 | 187 | 159 | 132 | 189 | 160 | 133 | 190 | 160 | 134 | 190 | 1 | 1 | 3
D | 9 | D9 | 2279 | 1614 | 1614 | 1 | 0 | 0 | 143.5056 | 143 | 0.85584 | 143 | 143 | 143 | 144 | 145 | 1 | 42.0861 | 42.5 | 1.3997 | 41 | 41 | 42.5 | 43 | 43.5 | 2 | 58.3726 | 58.4314 | 0.44401 | 58.0392 | 58.0392 | 58.4314 | 58.8235 | 58.8235 | 0.78432 | 21.6326 | 22 | 0.84771 | 21 | 21 | 22 | 22 | 23 | 1 | -25.7862 | -26 | 0.85577 | -27 | -26 | -26 | -25 | -25 | 1 | 0.11462 | 145.2378 | 1.2378 | 0.14188 | 0 |  | 156.4461 | 128.9703 | 185.6667 | 157 | 129 | 186 | 1.2296 | 1.1838 | 1.2654 | 155 | 128 | 184 | 156 | 128 | 185 | 157 | 129 | 186 | 157 | 130 | 186 | 158 | 130 | 187 | 1 | 2 | 1
D | 10 | D10 | 2267 | 1606 | 1606 | 1 | 0 | 0 | 141.279 | 141 | 0.88709 | 140 | 141 | 141 | 142 | 142 | 1 | 42.5324 | 43 | 1.381 | 41 | 41 | 43 | 43.5 | 43.5 | 2.5 | 57.5921 | 57.6471 | 0.36435 | 57.2549 | 57.2549 | 57.6471 | 57.6471 | 58.0392 | 0.39216 | 21.8624 | 22 | 0.76227 | 21 | 21 | 22 | 22 | 23 | 1 | -25.8082 | -26 | 0.9356 | -27 | -26 | -26 | -25 | -25 | 1 | 0.077833 | 143.264 | 1.264 | 0.098381 | 0 |  | 154.655 | 126.5679 | 183.5455 | 155 | 126 | 184 | 0.89125 | 1.1825 | 1.5387 | 154 | 125 | 182 | 154 | 126 | 183 | 155 | 126 | 184 | 155 | 127 | 184 | 156 | 128 | 185 | 1 | 1 | 1
D | 11 | D11 | 2223 | 1568 | 1568 | 1 | 0 | 0 | 138.9981 | 139 | 0.87107 | 138 | 139 | 139 | 139 | 140 | 0 | 42.7736 | 43.5 | 1.4647 | 41 | 42.5 | 43.5 | 43.5 | 43.5 | 1 | 56.7014 | 56.8627 | 0.36449 | 56.4706 | 56.4706 | 56.8627 | 56.8627 | 57.2549 | 0.39216 | 22.2972 | 23 | 0.93691 | 21 | 22 | 23 | 23 | 23 | 1 | -25.9292 | -26 | 0.9189 | -27 | -27 | -26 | -25 | -25 | 2 | 0.066964 | 141.2381 | 1.2381 | 0.082908 | 0 |  | 152.4783 | 124.1614 | 181.3916 | 153 | 124 | 182 | 0.95971 | 1.1599 | 1.3882 | 151 | 123 | 180 | 152 | 123 | 181 | 153 | 124 | 182 | 153 | 125 | 182 | 153 | 125 | 183 | 1 | 2 | 1
D | 12 | D12 | 2165 | 1522 | 1522 | 1 | 0 | 0 | 138.1209 | 138 | 1.3278 | 137 | 137 | 138 | 139 | 140 | 2 | 39.6754 | 40 | 1.9259 | 37 | 38 | 40 | 41 | 41 | 3 | 56.4492 | 56.4706 | 0.50542 | 56.0784 | 56.0784 | 56.4706 | 56.8627 | 57.2549 | 0.78431 | 20.5657 | 21 | 0.93789 | 19 | 20 | 21 | 21 | 21 | 1 | -24.2562 | -24 | 1.1604 | -25 | -25 | -24 | -24 | -23 | 1 | 0.032852 | 141.7 | 1.7 | 0.055848 | 0 |  | 150.5256 | 124.3666 | 177.5585 | 151 | 124 | 178 | 1.4307 | 1.6122 | 2.0244 | 149 | 123 | 175 | 150 | 123 | 176 | 151 | 124 | 178 | 151 | 125 | 179 | 152 | 126 | 180 | 1 | 2 | 3
E | 1 | E1 | 1669 | 1091 | 1091 | 1 | 0 | 0 | 208.2786 | 208 | 0.87225 | 207 | 208 | 208 | 209 | 209 | 1 | -2.104 | -2.5 | 0.88538 | -3.5 | -2.5 | -2.5 | -1 | -1 | 1.5 | 83.6477 | 83.5294 | 0.32307 | 83.1373 | 83.5294 | 83.5294 | 83.9216 | 83.9216 | 0.39216 | -0.73327 | -1 | 0.62176 | -1 | -1 | -1 | 0 | 0 | 1 | 1.1907 | 1 | 0.39299 | 1 | 1 | 1 | 1 | 2 | 0 | 0.012832 | 210 | 1 | 0.012832 | 0 |  | 207.7278 | 208.9166 | 205.8973 | 208 | 209 | 206 | 1.0537 | 0.91241 | 1.3581 | 206 | 208 | 204 | 207 | 209 | 205 | 208 | 209 | 206 | 209 | 209 | 207 | 209 | 210 | 207 | 2 | 0 | 2
E | 2 | E2 | 1726 | 1141 | 1141 | 1 | 0 | 0 | 154.0114 | 154 | 1.1077 | 153 | 153 | 154 | 155 | 156 | 2 | 25.6713 | 25 | 1.2948 | 24 | 25 | 25 | 27 | 27 | 2 | 62.753 | 62.7451 | 0.43435 | 62.3529 | 62.3529 | 62.7451 | 63.1373 | 63.5294 | 0.78431 | 12.4505 | 12 | 0.62999 | 12 | 12 | 12 | 13 | 13 | 1 | -17.8247 | -18 | 0.5826 | -18 | -18 | -18 | -18 | -17 | 0 | 0.12007 | 156.1752 | 1.1752 | 0.1411 | 0 |  | 158.929 | 145.631 | 183.6757 | 159 | 146 | 184 | 1.313 | 1.2772 | 1.6512 | 157 | 144 | 181 | 158 | 145 | 183 | 159 | 146 | 184 | 160 | 146 | 185 | 161 | 147 | 186 | 2 | 1 | 2
E | 3 | E3 | 1815 | 1214 | 1214 | 1 | 0 | 0 | 146.3007 | 146 | 0.81072 | 146 | 146 | 146 | 147 | 147 | 1 | 28.7912 | 29 | 1.4551 | 27 | 28 | 29 | 30.5 | 30.5 | 2.5 | 59.7597 | 59.6078 | 0.34264 | 59.6078 | 59.6078 | 59.6078 | 60 | 60.3922 | 0.39215 | 14.2158 | 14 | 0.75173 | 13 | 14 | 14 | 15 | 15 | 1 | -20.0857 | -21 | 1.173 | -21 | -21 | -21 | -19 | -18 | 2 | 0.079077 | 148.3333 | 1.3333 | 0.10544 | 0 |  | 152.3657 | 137.0239 | 179.2644 | 152 | 137 | 180 | 1.0303 | 1.1062 | 1.5079 | 151 | 136 | 177 | 152 | 136 | 178 | 152 | 137 | 180 | 153 | 137 | 180 | 153 | 138 | 181 | 1 | 1 | 2
E | 4 | E4 | 1887 | 1282 | 1282 | 1 | 0 | 0 | 142.6115 | 143 | 0.82721 | 142 | 142 | 143 | 143 | 144 | 1 | 31.5749 | 31 | 1.1546 | 30.5 | 31 | 31 | 32.5 | 32.5 | 1.5 | 58.2754 | 58.4314 | 0.33506 | 58.0392 | 58.0392 | 58.4314 | 58.4314 | 58.8235 | 0.39216 | 15.7785 | 16 | 0.71402 | 15 | 15 | 16 | 16 | 16 | 1 | -21.2894 | -21 | 0.66801 | -22 | -21 | -21 | -21 | -21 | 0 | 0.11778 | 144.2119 | 1.2119 | 0.14275 | 0 |  | 150.0078 | 132.163 | 177.468 | 150 | 132 | 177 | 0.99488 | 1.0532 | 1.285 | 149 | 131 | 176 | 149 | 132 | 177 | 150 | 132 | 177 | 151 | 133 | 178 | 151 | 133 | 179 | 2 | 1 | 1
E | 5 | E5 | 1960 | 1340 | 1340 | 1 | 0 | 0 | 139.9776 | 140 | 0.78605 | 139 | 140 | 140 | 140 | 141 | 0 | 34.141 | 34.5 | 1.4844 | 32.5 | 32.5 | 34.5 | 35.5 | 35.5 | 3 | 57.1428 | 57.2549 | 0.3764 | 56.4706 | 56.8627 | 57.2549 | 57.2549 | 57.6471 | 0.39215 | 17.5657 | 18 | 0.73999 | 17 | 17 | 18 | 18 | 18 | 1 | -23.1552 | -23 | 1.0391 | -24 | -24 | -23 | -22 | -22 | 2 | 0.028358 | 142.0789 | 1.0789 | 0.030597 | 0 |  | 147.7903 | 128.441 | 177.3739 | 148 | 128 | 177 | 0.79678 | 1.0557 | 2.1594 | 147 | 127 | 175 | 147 | 128 | 176 | 148 | 128 | 177 | 148 | 129 | 179 | 149 | 130 | 180 | 1 | 1 | 3
E | 6 | E6 | 2017 | 1390 | 1390 | 1 | 0 | 0 | 137.618 | 137 | 0.89642 | 137 | 137 | 137 | 138 | 139 | 1 | 37.4748 | 37.5 | 1.1348 | 35.5 | 37.5 | 37.5 | 38 | 38.5 | 0.5 | 56.171 | 56.0784 | 0.40488 | 55.6863 | 56.0784 | 56.0784 | 56.4706 | 56.8627 | 0.39215 | 19.3288 | 19 | 0.76869 | 18 | 19 | 19 | 20 | 20 | 1 | -23.9403 | -24 | 0.51893 | -24 | -24 | -24 | -24 | -23 | 0 | 0.020863 | 140.2759 | 1.2759 | 0.026619 | 0 |  | 147.8885 | 124.7151 | 176.4914 | 148 | 125 | 176 | 1.377 | 1.0396 | 1.195 | 146 | 124 | 175 | 147 | 124 | 176 | 148 | 125 | 176 | 149 | 125 | 177 | 150 | 126 | 178 | 2 | 1 | 1
E | 7 | E7 | 2059 | 1427 | 1427 | 1 | 0 | 0 | 136.0084 | 136 | 0.88936 | 135 | 136 | 136 | 136 | 137 | 0 | 38.9902 | 39.5 | 1.3431 | 37.5 | 38 | 39.5 | 40 | 40 | 2 | 55.5167 | 55.2941 | 0.36954 | 55.2941 | 55.2941 | 55.2941 | 55.6863 | 56.0784 | 0.39216 | 20.2845 | 20 | 0.81728 | 19 | 20 | 20 | 21 | 21 | 1 | -25.0259 | -25 | 1.1135 | -26 | -26 | -25 | -24 | -24 | 2 | 0.035739 | 138.1765 | 1.1765 | 0.042046 | 0 |  | 146.8066 | 122.5501 | 176.274 | 147 | 122 | 177 | 0.98752 | 0.987 | 1.946 | 146 | 122 | 174 | 146 | 122 | 175 | 147 | 122 | 177 | 147 | 123 | 177 | 148 | 124 | 179 | 1 | 1 | 2
E | 8 | E8 | 2088 | 1452 | 1452 | 1 | 0 | 0 | 133.1784 | 133 | 1.066 | 132 | 133 | 133 | 134 | 135 | 1 | 40.5121 | 41 | 1.5356 | 38.5 | 40 | 41 | 41 | 43 | 1 | 54.4796 | 54.5098 | 0.43791 | 54.1176 | 54.1176 | 54.5098 | 54.5098 | 54.902 | 0.39215 | 20.9972 | 21 | 0.95811 | 20 | 21 | 21 | 21 | 23 | 0 | -26.3409 | -27 | 0.9618 | -27 | -27 | -27 | -26 | -25 | 1 | 0.10193 | 135.527 | 1.527 | 0.15565 | 0 |  | 144.2135 | 119.292 | 175.3946 | 144 | 119 | 176 | 1.1886 | 1.4307 | 1.5009 | 143 | 118 | 174 | 143 | 118 | 174 | 144 | 119 | 176 | 145 | 120 | 176 | 146 | 121 | 177 | 2 | 2 | 2
E | 9 | E9 | 2100 | 1461 | 1461 | 1 | 0 | 0 | 131.2444 | 131 | 0.77006 | 130 | 131 | 131 | 132 | 132 | 1 | 42.1725 | 43 | 1.4164 | 40 | 41 | 43 | 43 | 43 | 2 | 53.5916 | 53.3333 | 0.42149 | 53.3333 | 53.3333 | 53.3333 | 54.1176 | 54.1176 | 0.78431 | 22.3778 | 23 | 0.91914 | 21 | 22 | 23 | 23 | 23 | 1 | -26.4894 | -27 | 0.94474 | -27 | -27 | -27 | -26 | -25 | 1 | 0.05065 | 133.2027 | 1.2027 | 0.060917 | 0 |  | 143.9035 | 116.616 | 173.6735 | 144 | 116 | 174 | 1.1139 | 1.0827 | 1.3015 | 142 | 116 | 172 | 144 | 116 | 173 | 144 | 116 | 174 | 145 | 117 | 174 | 145 | 118 | 175 | 1 | 1 | 1
E | 10 | E10 | 2087 | 1449 | 1449 | 1 | 0 | 0 | 129.5273 | 129 | 0.9303 | 129 | 129 | 129 | 130 | 131 | 1 | 41.8847 | 43 | 1.5292 | 40 | 41 | 43 | 43 | 43 | 2 | 53.067 | 52.9412 | 0.3399 | 52.9412 | 52.9412 | 52.9412 | 53.3333 | 53.7255 | 0.39216 | 21.7895 | 22 | 0.68468 | 21 | 21 | 22 | 22 | 22 | 1 | -26.2761 | -27 | 1.0641 | -27 | -27 | -27 | -25 | -25 | 2 | 0.040028 | 132.0345 | 1.0345 | 0.041408 | 0 |  | 142.1063 | 115.0028 | 171.6687 | 142 | 115 | 172 | 1.0136 | 1.2933 | 1.4733 | 141 | 114 | 170 | 142 | 114 | 171 | 142 | 115 | 172 | 143 | 116 | 172 | 143 | 117 | 173 | 1 | 2 | 1
E | 11 | E11 | 2049 | 1419 | 1419 | 1 | 0 | 0 | 128.6342 | 129 | 0.84862 | 128 | 128 | 129 | 129 | 130 | 1 | 40.3971 | 41 | 1.5908 | 38 | 40 | 41 | 41 | 41.5 | 1 | 52.7867 | 52.9412 | 0.32381 | 52.549 | 52.549 | 52.9412 | 52.9412 | 52.9412 | 0.39216 | 21.4123 | 22 | 0.98573 | 20 | 21 | 22 | 22 | 22 | 1 | -24.685 | -25 | 0.88324 | -26 | -25 | -25 | -24 | -24 | 1 | 0.11416 | 130.3642 | 1.3642 | 0.15574 | 0 |  | 141.4299 | 114.6885 | 168.7414 | 141 | 114 | 169 | 0.96912 | 1.1758 | 1.4791 | 140 | 114 | 167 | 141 | 114 | 168 | 141 | 114 | 169 | 142 | 115 | 170 | 143 | 116 | 170 | 1 | 1 | 2
E | 12 | E12 | 1994 | 1371 | 1371 | 1 | 0 | 0 | 129.9059 | 130 | 1.1369 | 129 | 129 | 130 | 131 | 131 | 2 | 34.9369 | 35.5 | 1.6866 | 32.5 | 34.5 | 35.5 | 35.5 | 37 | 1 | 53.2976 | 53.3333 | 0.44996 | 52.9412 | 52.9412 | 53.3333 | 53.7255 | 53.7255 | 0.78431 | 18.2035 | 18 | 0.92658 | 17 | 18 | 18 | 19 | 19 | 1 | -22.515 | -23 | 0.74796 | -23 | -23 | -23 | -22 | -22 | 1 | 0.081692 | 132.4286 | 1.4286 | 0.1167 | 0 |  | 139.6878 | 117.9059 | 165.9978 | 140 | 118 | 166 | 1.0074 | 1.6068 | 1.3884 | 139 | 116 | 164 | 139 | 117 | 165 | 140 | 118 | 166 | 140 | 119 | 167 | 141 | 120 | 168 | 1 | 2 | 2
F | 1 | F1 | 1577 | 1013 | 1013 | 1 | 0 | 0 | 209.3889 | 210 | 0.96723 | 208 | 209 | 210 | 210 | 210 | 1 | -2.2744 | -3 | 0.99366 | -3 | -3 | -3 | -1 | -1 | 2 | 84.0919 | 84.3137 | 0.33463 | 83.5294 | 83.9216 | 84.3137 | 84.3137 | 84.3137 | 0.39216 | -0.62488 | -1 | 0.48439 | -1 | -1 | -1 | 0 | 0 | 1 | 1.6742 | 2 | 0.56451 | 1 | 1 | 2 | 2 | 2 | 1 | 0.043435 | 211.0455 | 1.0455 | 0.04541 | 0 |  | 209.3889 | 210.0138 | 206.0898 | 210 | 210 | 206 | 0.96723 | 0.90005 | 1.6457 | 208 | 209 | 204 | 209 | 210 | 205 | 210 | 210 | 206 | 210 | 211 | 208 | 210 | 211 | 208 | 1 | 1 | 3
F | 2 | F2 | 1626 | 1056 | 1056 | 1 | 0 | 0 | 153.2178 | 153 | 0.81799 | 152 | 153 | 153 | 154 | 154 | 1 | 25.5592 | 25 | 1.5679 | 24 | 24 | 25 | 27 | 28 | 3 | 62.4584 | 62.3529 | 0.31798 | 61.9608 | 62.3529 | 62.3529 | 62.7451 | 62.7451 | 0.39215 | 12.4697 | 12 | 0.766 | 12 | 12 | 12 | 13 | 14 | 1 | -17.6809 | -18 | 0.74875 | -19 | -18 | -18 | -17 | -17 | 1 | 0.078598 | 155.1446 | 1.1446 | 0.089962 | 0 |  | 158.1979 | 144.8504 | 182.6212 | 158 | 145 | 183 | 1.0502 | 1.133 | 1.4023 | 157 | 144 | 181 | 157 | 144 | 182 | 158 | 145 | 183 | 159 | 145 | 183 | 159 | 146 | 185 | 2 | 1 | 1
F | 3 | F3 | 1695 | 1114 | 1114 | 1 | 0 | 0 | 146.7433 | 147 | 0.79564 | 146 | 146 | 147 | 147 | 148 | 1 | 29.1005 | 29.5 | 1.2604 | 27.5 | 28 | 29.5 | 30.5 | 30.5 | 2.5 | 59.9321 | 60 | 0.34221 | 59.6078 | 59.6078 | 60 | 60 | 60.3922 | 0.39215 | 14.5171 | 15 | 0.62601 | 14 | 14 | 15 | 15 | 15 | 1 | -19.8474 | -20 | 1.0392 | -21 | -21 | -20 | -19 | -18 | 2 | 0.018851 | 149.0952 | 1.0952 | 0.020646 | 0 |  | 153.2657 | 137.237 | 179.4093 | 153 | 137 | 179 | 0.84001 | 1.0896 | 1.6558 | 152 | 136 | 177 | 153 | 136 | 178 | 153 | 137 | 179 | 154 | 138 | 181 | 154 | 139 | 181 | 1 | 2 | 3
F | 4 | F4 | 1756 | 1166 | 1166 | 1 | 0 | 0 | 142.7316 | 143 | 1.1486 | 141 | 142 | 143 | 143 | 144 | 1 | 31.2834 | 31 | 0.88699 | 30.5 | 31 | 31 | 31 | 32.5 | 0 | 58.3244 | 58.4314 | 0.46376 | 57.6471 | 58.0392 | 58.4314 | 58.4314 | 58.8235 | 0.39216 | 15.7101 | 16 | 0.63247 | 15 | 15 | 16 | 16 | 16 | 1 | -21.205 | -21 | 0.5559 | -22 | -21 | -21 | -21 | -21 | 0 | 0.081475 | 145.3053 | 1.3053 | 0.10635 | 0 |  | 149.9039 | 132.4906 | 177.6441 | 150 | 132 | 177 | 1.2309 | 1.2359 | 1.2585 | 149 | 131 | 176 | 149 | 132 | 177 | 150 | 132 | 177 | 150 | 133 | 178 | 151 | 134 | 180 | 1 | 1 | 1
F | 5 | F5 | 1807 | 1211 | 1211 | 1 | 0 | 0 | 139.4649 | 139 | 0.88666 | 139 | 139 | 139 | 140 | 141 | 1 | 34.4344 | 34.5 | 1.3989 | 32.5 | 34.5 | 34.5 | 35.5 | 35.5 | 1 | 56.8725 | 56.8627 | 0.44631 | 56.4706 | 56.4706 | 56.8627 | 57.2549 | 57.2549 | 0.78431 | 17.7614 | 18 | 0.73182 | 17 | 18 | 18 | 18 | 18 | 0 | -23.2469 | -23 | 0.99883 | -25 | -24 | -23 | -23 | -22 | 1 | 0.1057 | 141.2656 | 1.2656 | 0.13377 | 0 |  | 147.3287 | 127.7539 | 177.0479 | 147 | 127 | 176 | 0.81474 | 1.2008 | 2.0488 | 146 | 127 | 175 | 147 | 127 | 176 | 147 | 127 | 176 | 148 | 128 | 178 | 148 | 129 | 180 | 1 | 1 | 2
F | 6 | F6 | 1859 | 1257 | 1257 | 1 | 0 | 0 | 136.9181 | 137 | 0.94623 | 136 | 136 | 137 | 137 | 138 | 1 | 36.5243 | 37.5 | 1.4131 | 34.5 | 35.5 | 37.5 | 37.5 | 38.5 | 2 | 55.8803 | 55.6863 | 0.38079 | 55.6863 | 55.6863 | 55.6863 | 56.0784 | 56.4706 | 0.39216 | 18.6348 | 19 | 0.70789 | 18 | 18 | 19 | 19 | 20 | 1 | -23.7454 | -24 | 0.78224 | -25 | -24 | -24 | -23 | -23 | 1 | 0.038982 | 139.3061 | 1.3061 | 0.050915 | 0 |  | 146.4471 | 124.4033 | 175.4081 | 147 | 124 | 175 | 1.1907 | 1.0724 | 1.9715 | 145 | 123 | 173 | 146 | 124 | 174 | 147 | 124 | 175 | 147 | 125 | 177 | 148 | 126 | 178 | 1 | 1 | 3
F | 7 | F7 | 1901 | 1289 | 1289 | 1 | 0 | 0 | 135.3631 | 135 | 0.83781 | 134 | 135 | 135 | 136 | 136 | 1 | 38.0582 | 38 | 1.2427 | 36.5 | 37.5 | 38 | 38.5 | 40 | 1 | 55.3808 | 55.2941 | 0.36732 | 54.902 | 55.2941 | 55.2941 | 55.6863 | 55.6863 | 0.39216 | 19.7122 | 20 | 0.81628 | 19 | 19 | 20 | 20 | 21 | 1 | -24.4383 | -24 | 0.97458 | -26 | -25 | -24 | -24 | -23 | 1 | 0.06827 | 137.2273 | 1.2273 | 0.083786 | 0 |  | 145.7967 | 122.3351 | 174.9899 | 146 | 122 | 175 | 0.78848 | 1.0615 | 1.8167 | 145 | 121 | 173 | 145 | 122 | 174 | 146 | 122 | 175 | 146 | 123 | 176 | 147 | 124 | 177 | 1 | 1 | 2
F | 8 | F8 | 1932 | 1316 | 1316 | 1 | 0 | 0 | 133.2386 | 133 | 0.98316 | 132 | 133 | 133 | 134 | 134 | 1 | 40.7128 | 41 | 1.5546 | 38 | 40 | 41 | 41 | 43 | 1 | 54.4356 | 54.5098 | 0.38689 | 54.1176 | 54.1176 | 54.5098 | 54.5098 | 54.902 | 0.39215 | 21.3435 | 21 | 1.0112 | 20 | 21 | 21 | 22 | 23 | 1 | -25.5859 | -25 | 1.0389 | -27 | -26 | -25 | -25 | -24 | 1 | 0.069149 | 135.2747 | 1.2747 | 0.088146 | 0 |  | 145.329 | 119.1892 | 174.4749 | 145 | 119 | 174 | 1.1971 | 1.1956 | 1.691 | 144 | 118 | 172 | 144 | 118 | 173 | 145 | 119 | 174 | 146 | 120 | 176 | 147 | 121 | 177 | 2 | 2 | 3
F | 9 | F9 | 1936 | 1319 | 1319 | 1 | 0 | 0 | 130.2578 | 130 | 0.9162 | 129 | 130 | 130 | 131 | 131 | 1 | 41.6202 | 41.5 | 1.7402 | 40 | 41 | 41.5 | 43 | 43 | 2 | 53.2655 | 53.3333 | 0.39932 | 52.9412 | 52.9412 | 53.3333 | 53.3333 | 53.7255 | 0.39216 | 21.8044 | 22 | 0.90335 | 21 | 21 | 22 | 22 | 23 | 1 | -26.0842 | -27 | 1.1607 | -27 | -27 | -27 | -25 | -24 | 2 | 0.072782 | 132.2188 | 1.2188 | 0.088704 | 0 |  | 142.7635 | 115.8484 | 172.1736 | 143 | 116 | 172 | 1.1976 | 1.267 | 1.6983 | 141 | 115 | 170 | 142 | 115 | 171 | 143 | 116 | 172 | 143 | 117 | 173 | 144 | 117 | 174 | 1 | 2 | 2
F | 10 | F10 | 1925 | 1311 | 1311 | 1 | 0 | 0 | 128.8085 | 129 | 1.0371 | 128 | 128 | 129 | 129 | 130 | 1 | 41.0721 | 41 | 1.9798 | 38 | 40 | 41 | 43 | 43 | 3 | 52.7635 | 52.9412 | 0.41128 | 52.1569 | 52.549 | 52.9412 | 52.9412 | 53.3333 | 0.39216 | 21.4905 | 22 | 0.93435 | 20 | 21 | 22 | 22 | 22 | 1 | -25.8825 | -26 | 1.2056 | -27 | -27 | -26 | -25 | -24 | 2 | 0.050343 | 131.197 | 1.197 | 0.060259 | 0 |  | 141.0473 | 114.5782 | 170.2532 | 141 | 114 | 170 | 1.3288 | 1.4179 | 1.8033 | 139 | 113 | 168 | 140 | 114 | 169 | 141 | 114 | 170 | 142 | 115 | 172 | 143 | 116 | 172 | 2 | 1 | 3
F | 11 | F11 | 1890 | 1282 | 1282 | 1 | 0 | 0 | 128.1864 | 128 | 0.89252 | 127 | 128 | 128 | 129 | 129 | 1 | 39.2531 | 40 | 1.8105 | 37 | 38 | 40 | 40 | 41 | 2 | 52.5236 | 52.549 | 0.39408 | 52.1569 | 52.1569 | 52.549 | 52.9412 | 52.9412 | 0.78431 | 20.6919 | 21 | 1.0695 | 19 | 20 | 21 | 21 | 22 | 1 | -24.2566 | -24 | 1.264 | -26 | -25 | -24 | -23 | -23 | 2 | 0.070203 | 130.3111 | 1.3111 | 0.092044 | 0 |  | 140.248 | 114.4813 | 167.2207 | 140 | 114 | 167 | 1.2341 | 1.2694 | 1.9463 | 139 | 113 | 165 | 139 | 114 | 166 | 140 | 114 | 167 | 141 | 115 | 169 | 142 | 116 | 170 | 2 | 1 | 3
F | 12 | F12 | 1852 | 1249 | 1249 | 1 | 0 | 0 | 130.7398 | 130 | 1.2818 | 130 | 130 | 130 | 131 | 133 | 1 | 34.8351 | 34.5 | 1.9003 | 32.5 | 34.5 | 34.5 | 36 | 37 | 1.5 | 53.5892 | 53.3333 | 0.54237 | 52.9412 | 53.3333 | 53.3333 | 53.7255 | 54.5098 | 0.39215 | 18.2074 | 18 | 1.0336 | 17 | 18 | 18 | 19 | 19 | 1 | -22.293 | -22 | 0.90704 | -23 | -23 | -22 | -22 | -21 | 1 | 0.11049 | 133.6377 | 1.6377 | 0.18094 | 0 |  | 140.5428 | 118.6653 | 166.458 | 141 | 118 | 167 | 1.0801 | 1.8691 | 1.5953 | 139 | 117 | 165 | 140 | 117 | 165 | 141 | 118 | 167 | 141 | 119 | 167 | 142 | 121 | 168 | 1 | 2 | 2
G | 1 | G1 | 1483 | 937 | 937 | 1 | 0 | 0 | 207.793 | 208 | 1.0928 | 207 | 207 | 208 | 208 | 209 | 1 | -2.3799 | -3 | 0.94125 | -3 | -3 | -3 | -1 | -1 | 2 | 83.5307 | 83.5294 | 0.34466 | 83.1373 | 83.5294 | 83.5294 | 83.5294 | 83.9216 | 0 | -0.6841 | -1 | 0.46512 | -1 | -1 | -1 | 0 | 0 | 1 | 1.7076 | 2 | 0.50412 | 1 | 1 | 2 | 2 | 2 | 1 | 0.026681 | 210.2 | 1.2 | 0.032017 | 0 |  | 207.793 | 208.4771 | 204.4013 | 208 | 208 | 204 | 1.0928 | 1.1537 | 1.4911 | 207 | 207 | 203 | 207 | 208 | 204 | 208 | 208 | 204 | 208 | 209 | 206 | 209 | 210 | 206 | 1 | 1 | 2
G | 2 | G2 | 1530 | 977 | 977 | 1 | 0 | 0 | 152.5097 | 152 | 0.84213 | 152 | 152 | 152 | 153 | 154 | 1 | 25.2042 | 25 | 1.2028 | 24 | 25 | 25 | 26 | 27 | 1 | 62.1956 | 61.9608 | 0.34439 | 61.9608 | 61.9608 | 61.9608 | 62.3529 | 62.7451 | 0.39216 | 12.2364 | 12 | 0.64368 | 12 | 12 | 12 | 12 | 13 | 0 | -17.7462 | -18 | 0.70787 | -18 | -18 | -18 | -17 | -17 | 1 | 0.10235 | 154.27 | 1.27 | 0.12999 | 0 |  | 157.0809 | 144.3582 | 182.044 | 157 | 144 | 182 | 1.122 | 1.081 | 1.3883 | 156 | 143 | 180 | 156 | 144 | 181 | 157 | 144 | 182 | 158 | 145 | 183 | 159 | 146 | 184 | 2 | 1 | 2
G | 3 | G3 | 1578 | 1016 | 1016 | 1 | 0 | 0 | 147.1585 | 147 | 0.73175 | 146 | 147 | 147 | 147 | 148 | 0 | 28.1949 | 28 | 1.1767 | 27 | 28 | 28 | 29.5 | 29.5 | 1.5 | 60.0957 | 60 | 0.32508 | 59.6078 | 60 | 60 | 60.3922 | 60.3922 | 0.39216 | 14.1348 | 14 | 0.72161 | 13 | 14 | 14 | 15 | 15 | 1 | -19.1506 | -19 | 0.87885 | -20 | -20 | -19 | -19 | -18 | 1 | 0.051181 | 149.2115 | 1.2115 | 0.062008 | 0 |  | 153.4026 | 137.9321 | 178.8514 | 153 | 138 | 179 | 0.85866 | 1.025 | 1.2906 | 152 | 137 | 177 | 153 | 137 | 179 | 153 | 138 | 179 | 154 | 138 | 179 | 154 | 139 | 180 | 1 | 1 | 0
G | 4 | G4 | 1637 | 1064 | 1064 | 1 | 0 | 0 | 142.5489 | 142 | 0.86831 | 142 | 142 | 142 | 143 | 144 | 1 | 30.7773 | 30.5 | 1.1403 | 29.5 | 30.5 | 30.5 | 31 | 32.5 | 0.5 | 58.2648 | 58.0392 | 0.35314 | 58.0392 | 58.0392 | 58.0392 | 58.4314 | 58.8235 | 0.39216 | 15.4135 | 15 | 0.67897 | 15 | 15 | 15 | 16 | 16 | 1 | -21.0273 | -21 | 0.69686 | -22 | -21 | -21 | -21 | -20 | 0 | 0.030075 | 145.0312 | 1.0312 | 0.031015 | 0 |  | 149.6034 | 132.4558 | 176.8628 | 150 | 132 | 177 | 0.89033 | 1.0973 | 1.2291 | 149 | 131 | 175 | 149 | 132 | 176 | 150 | 132 | 177 | 150 | 133 | 177 | 151 | 134 | 178 | 1 | 1 | 1
G | 5 | G5 | 1690 | 1109 | 1109 | 1 | 0 | 0 | 138.6528 | 139 | 0.9516 | 138 | 138 | 139 | 139 | 140 | 1 | 33.2187 | 33 | 1.6127 | 31 | 32 | 33 | 34.5 | 34.5 | 2.5 | 56.5756 | 56.4706 | 0.46969 | 56.0784 | 56.0784 | 56.4706 | 56.8627 | 57.2549 | 0.78431 | 17.1668 | 17 | 0.93476 | 16 | 16 | 17 | 18 | 18 | 2 | -22.3832 | -23 | 0.89024 | -23 | -23 | -23 | -21 | -21 | 2 | 0.030658 | 141.2059 | 1.2059 | 0.03697 | 0 |  | 146.4238 | 127.4211 | 174.8557 | 146 | 127 | 175 | 1.2162 | 1.3963 | 1.3222 | 145 | 126 | 173 | 146 | 126 | 174 | 146 | 127 | 175 | 147 | 128 | 176 | 148 | 129 | 176 | 1 | 2 | 2
G | 6 | G6 | 1736 | 1150 | 1150 | 1 | 0 | 0 | 136.8643 | 137 | 0.88052 | 136 | 136 | 137 | 137 | 138 | 1 | 34.9935 | 34.5 | 1.6187 | 33 | 34.5 | 34.5 | 35.5 | 37.5 | 1 | 55.8271 | 55.6863 | 0.39108 | 55.2941 | 55.6863 | 55.6863 | 56.0784 | 56.4706 | 0.39216 | 18.0548 | 18 | 0.75764 | 17 | 18 | 18 | 18 | 19 | 0 | -23.373 | -23 | 0.88157 | -24 | -24 | -23 | -23 | -23 | 1 | 0.01913 | 139.3636 | 1.3636 | 0.026087 | 0 |  | 145.113 | 124.8887 | 174.6513 | 145 | 125 | 175 | 1.0796 | 1.1636 | 1.8255 | 144 | 124 | 173 | 144 | 124 | 173 | 145 | 125 | 175 | 146 | 125 | 176 | 147 | 126 | 177 | 2 | 1 | 3
G | 7 | G7 | 1768 | 1178 | 1178 | 1 | 0 | 0 | 135.0526 | 135 | 0.70364 | 134 | 135 | 135 | 135 | 136 | 0 | 36.8468 | 37.5 | 1.399 | 35.5 | 35.5 | 37.5 | 37.5 | 37.5 | 2 | 55.2702 | 55.2941 | 0.34284 | 54.5098 | 55.2941 | 55.2941 | 55.2941 | 55.6863 | 0 | 18.7479 | 19 | 0.68861 | 18 | 18 | 19 | 19 | 19 | 1 | -23.7419 | -24 | 0.74059 | -24 | -24 | -24 | -24 | -23 | 0 | 0.036503 | 137.2558 | 1.2558 | 0.04584 | 0 |  | 144.8896 | 122.3625 | 173.5289 | 145 | 122 | 174 | 0.74065 | 1.0647 | 1.38 | 144 | 121 | 172 | 145 | 122 | 173 | 145 | 122 | 174 | 145 | 123 | 174 | 146 | 124 | 175 | 0 | 1 | 1
G | 8 | G8 | 1785 | 1190 | 1190 | 1 | 0 | 0 | 132.6034 | 132 | 0.90173 | 132 | 132 | 132 | 133 | 134 | 1 | 37.5 | 37.5 | 1.4086 | 35.5 | 37 | 37.5 | 38 | 40 | 1 | 54.3052 | 54.1176 | 0.347 | 54.1176 | 54.1176 | 54.1176 | 54.5098 | 54.5098 | 0.39215 | 19.4622 | 20 | 0.92521 | 18 | 19 | 20 | 20 | 21 | 1 | -23.9958 | -24 | 1.0008 | -25 | -24 | -24 | -23 | -23 | 1 | 0.042017 | 135.14 | 1.14 | 0.047899 | 0 |  | 142.9622 | 119.6899 | 171.4176 | 143 | 119 | 171 | 1.0035 | 1.1129 | 2.015 | 142 | 119 | 169 | 142 | 119 | 170 | 143 | 119 | 171 | 143 | 120 | 173 | 144 | 121 | 174 | 1 | 1 | 3
G | 9 | G9 | 1798 | 1203 | 1203 | 1 | 0 | 0 | 131.0141 | 131 | 0.84819 | 130 | 131 | 131 | 131 | 132 | 0 | 39.0794 | 40 | 1.5324 | 37 | 38 | 40 | 40 | 40 | 2 | 53.5165 | 53.3333 | 0.41118 | 52.9412 | 53.3333 | 53.3333 | 53.7255 | 54.1176 | 0.39215 | 20.5121 | 21 | 0.88929 | 19 | 20 | 21 | 21 | 21 | 1 | -24.9393 | -26 | 1.2965 | -26 | -26 | -26 | -24 | -23 | 2 | 0.035744 | 133.2326 | 1.2326 | 0.044057 | 0 |  | 142.1396 | 117.4165 | 170.852 | 142 | 117 | 171 | 1.0833 | 1.1852 | 1.7194 | 141 | 116 | 168 | 142 | 117 | 170 | 142 | 117 | 171 | 143 | 118 | 172 | 144 | 119 | 173 | 1 | 1 | 2
G | 10 | G10 | 1787 | 1195 | 1195 | 1 | 0 | 0 | 129.0209 | 129 | 0.78538 | 128 | 129 | 129 | 129 | 130 | 0 | 37.7205 | 38 | 1.3511 | 35.5 | 38 | 38 | 38 | 38.5 | 0 | 52.9133 | 52.9412 | 0.33223 | 52.549 | 52.9412 | 52.9412 | 52.9412 | 53.3333 | 0 | 19.7967 | 20 | 0.78169 | 19 | 20 | 20 | 20 | 20 | 0 | -23.9732 | -24 | 1.0204 | -26 | -24 | -24 | -24 | -23 | 0 | 0.037657 | 131.3111 | 1.3111 | 0.049372 | 0 |  | 139.8243 | 116.051 | 167.7188 | 140 | 116 | 168 | 0.86873 | 1.0442 | 1.4704 | 139 | 115 | 166 | 139 | 116 | 167 | 140 | 116 | 168 | 140 | 116 | 168 | 141 | 117 | 169 | 1 | 0 | 1
G | 11 | G11 | 1765 | 1173 | 1173 | 1 | 0 | 0 | 128.543 | 128 | 0.91679 | 128 | 128 | 128 | 129 | 130 | 1 | 38.0631 | 38 | 1.7705 | 36 | 37 | 38 | 40 | 40 | 3 | 52.663 | 52.549 | 0.38844 | 52.1569 | 52.549 | 52.549 | 52.9412 | 52.9412 | 0.39216 | 19.913 | 20 | 1.0307 | 19 | 19 | 20 | 21 | 21 | 2 | -24.1935 | -24 | 1.3351 | -26 | -26 | -24 | -23 | -23 | 3 | 0.11253 | 130.3788 | 1.3788 | 0.15516 | 0 |  | 139.4263 | 115.3359 | 167.3717 | 139 | 115 | 167 | 1.0265 | 1.1506 | 2.2383 | 138 | 114 | 165 | 139 | 115 | 166 | 139 | 115 | 167 | 140 | 116 | 169 | 141 | 117 | 170 | 1 | 1 | 3
G | 12 | G12 | 1725 | 1143 | 1143 | 1 | 0 | 0 | 129.657 | 129 | 1.5093 | 128 | 129 | 129 | 131 | 132 | 2 | 33.9274 | 34 | 1.3424 | 32 | 32.5 | 34 | 34.5 | 35.5 | 2 | 53.1704 | 52.9412 | 0.60428 | 52.549 | 52.549 | 52.9412 | 53.3333 | 54.1176 | 0.78432 | 17.769 | 18 | 0.77563 | 17 | 17 | 18 | 18 | 19 | 1 | -21.846 | -22 | 0.76418 | -23 | -22 | -22 | -22 | -21 | 0 | 0.048994 | 133.5179 | 1.5179 | 0.074366 | 0 |  | 139.1452 | 117.8731 | 164.4558 | 139 | 118 | 164 | 1.5364 | 1.7347 | 2.1251 | 137 | 116 | 162 | 138 | 117 | 163 | 139 | 118 | 164 | 140 | 119 | 166 | 141 | 120 | 167 | 2 | 2 | 3
H | 1 | H1 | 1396 | 862 | 862 | 1 | 0 | 0 | 207.0685 | 208 | 1.5328 | 205 | 206 | 208 | 208 | 208 | 2 | -2.7193 | -3 | 0.59621 | -3 | -3 | -3 | -3 | -1.5 | 0 | 83.3224 | 83.5294 | 0.4912 | 82.7451 | 83.1373 | 83.5294 | 83.5294 | 83.5294 | 0.39216 | -0.82019 | -1 | 0.38727 | -1 | -1 | -1 | -1 | 0 | 0 | 2.0441 | 2 | 0.2054 | 2 | 2 | 2 | 2 | 2 | 0 | 0.071926 | 209 | 1 | 0.071926 | 0 |  | 207.2703 | 207.8666 | 203.0244 | 208 | 208 | 204 | 1.6507 | 1.5156 | 1.5716 | 205 | 206 | 201 | 207 | 207 | 202 | 208 | 208 | 204 | 208 | 209 | 204 | 209 | 209 | 204 | 1 | 2 | 2
H | 2 | H2 | 1427 | 892 | 892 | 1 | 0 | 0 | 152.1076 | 152 | 1.53 | 150 | 152 | 152 | 153 | 154 | 1 | 23.0482 | 22.5 | 1.3627 | 21.5 | 22.5 | 22.5 | 24 | 25 | 1.5 | 62.1551 | 62.3529 | 0.56063 | 61.1765 | 61.9608 | 62.3529 | 62.3529 | 62.7451 | 0.39216 | 11.3061 | 11 | 0.76794 | 10 | 11 | 11 | 12 | 12 | 1 | -16.5179 | -17 | 0.89938 | -17 | -17 | -17 | -16 | -15 | 1 | 0.026906 | 155.1667 | 1.1667 | 0.03139 | 0 |  | 155.9731 | 144.7029 | 179.5291 | 156 | 145 | 180 | 1.8663 | 1.3443 | 2.4976 | 154 | 143 | 176 | 155 | 144 | 179 | 156 | 145 | 180 | 157 | 146 | 181 | 158 | 146 | 182 | 2 | 2 | 2
H | 3 | H3 | 1479 | 932 | 932 | 1 | 0 | 0 | 147.5601 | 148 | 1.2768 | 146 | 147 | 148 | 148 | 149 | 1 | 26.4394 | 27 | 1.5075 | 24 | 25 | 27 | 27 | 28 | 2 | 60.2958 | 60.3922 | 0.48852 | 59.6078 | 60 | 60.3922 | 60.7843 | 60.7843 | 0.78431 | 13.0815 | 13 | 0.85827 | 12 | 13 | 13 | 14 | 14 | 1 | -18.0075 | -18 | 0.82584 | -19 | -19 | -18 | -18 | -17 | 1 | 0.011803 | 150.0909 | 1.0909 | 0.012876 | 0 |  | 153.1416 | 138.9217 | 177.5805 | 154 | 139 | 178 | 1.7671 | 1.1293 | 2.4281 | 151 | 138 | 174 | 152 | 138 | 176 | 154 | 139 | 178 | 154 | 140 | 179 | 155 | 140 | 180 | 2 | 2 | 3
H | 4 | H4 | 1526 | 973 | 973 | 1 | 0 | 0 | 143.8099 | 144 | 0.91951 | 143 | 143 | 144 | 144 | 145 | 1 | 28.0216 | 28 | 1.4651 | 25 | 28 | 28 | 29 | 29 | 1 | 58.932 | 58.8235 | 0.42295 | 58.4314 | 58.4314 | 58.8235 | 59.2157 | 59.6078 | 0.78431 | 13.8633 | 14 | 0.62593 | 13 | 14 | 14 | 14 | 14 | 0 | -19.8469 | -20 | 1.3045 | -21 | -21 | -20 | -19 | -18 | 2 | 0.023638 | 146.2174 | 1.2174 | 0.028777 | 0 |  | 149.6228 | 134.9034 | 176.2271 | 150 | 135 | 177 | 0.9572 | 1.0138 | 2.3716 | 148 | 134 | 173 | 149 | 134 | 175 | 150 | 135 | 177 | 150 | 135 | 178 | 151 | 136 | 179 | 1 | 1 | 3
H | 5 | H5 | 1570 | 1007 | 1007 | 1 | 0 | 0 | 140.7478 | 141 | 0.8013 | 140 | 140 | 141 | 141 | 142 | 1 | 29.8952 | 30.5 | 1.2949 | 28 | 29.5 | 30.5 | 30.5 | 31 | 1 | 57.5844 | 57.6471 | 0.31261 | 57.2549 | 57.2549 | 57.6471 | 57.6471 | 58.0392 | 0.39216 | 14.9782 | 15 | 0.67438 | 14 | 15 | 15 | 15 | 16 | 0 | -20.4598 | -21 | 0.99807 | -21 | -21 | -21 | -20 | -20 | 1 | 0.10129 | 142.1275 | 1.1275 | 0.1142 | 0 |  | 147.6445 | 130.8937 | 173.9335 | 148 | 131 | 175 | 0.94833 | 0.87694 | 1.919 | 147 | 130 | 172 | 147 | 130 | 173 | 148 | 131 | 175 | 148 | 131 | 175 | 149 | 132 | 176 | 1 | 1 | 2
H | 6 | H6 | 1603 | 1036 | 1036 | 1 | 0 | 0 | 138.1448 | 138 | 0.98407 | 137 | 138 | 138 | 139 | 139 | 1 | 32.3036 | 32.5 | 1.4861 | 30.5 | 32 | 32.5 | 32.5 | 34.5 | 0.5 | 56.4729 | 56.4706 | 0.40794 | 56.0784 | 56.0784 | 56.4706 | 56.8627 | 56.8627 | 0.78431 | 16.6496 | 17 | 0.967 | 15 | 16 | 17 | 17 | 18 | 1 | -21.7461 | -22 | 1.0751 | -23 | -22 | -22 | -21 | -21 | 1 | 0.049228 | 140.1373 | 1.1373 | 0.055985 | 0 |  | 145.9431 | 127.1969 | 173.0579 | 146 | 127 | 173 | 0.9619 | 1.203 | 2.026 | 145 | 126 | 171 | 145 | 126 | 172 | 146 | 127 | 173 | 147 | 128 | 174 | 147 | 129 | 176 | 2 | 2 | 2
H | 7 | H7 | 1636 | 1064 | 1064 | 1 | 0 | 0 | 135.5789 | 136 | 0.91366 | 134 | 135 | 136 | 136 | 136 | 1 | 32.8947 | 32.5 | 1.7307 | 30.5 | 32 | 32.5 | 34.5 | 34.5 | 2.5 | 55.4216 | 55.2941 | 0.34348 | 54.902 | 55.2941 | 55.2941 | 55.6863 | 55.6863 | 0.39216 | 16.9652 | 17 | 1.0747 | 15 | 16 | 17 | 18 | 18 | 2 | -22.125 | -22 | 1.0529 | -23 | -23 | -22 | -21 | -21 | 2 | 0.091165 | 137.268 | 1.268 | 0.1156 | 0 |  | 143.3863 | 124.3919 | 171.187 | 144 | 124 | 171 | 1.1428 | 1.039 | 2.0806 | 142 | 123 | 169 | 143 | 124 | 170 | 144 | 124 | 171 | 144 | 125 | 173 | 144 | 126 | 174 | 1 | 1 | 3
H | 8 | H8 | 1652 | 1078 | 1078 | 1 | 0 | 0 | 134.8933 | 135 | 0.9043 | 134 | 134 | 135 | 135 | 136 | 1 | 33.0218 | 32.5 | 1.5197 | 31 | 32.5 | 32.5 | 34.5 | 34.5 | 2 | 55.2181 | 55.2941 | 0.3757 | 54.902 | 54.902 | 55.2941 | 55.2941 | 55.6863 | 0.39215 | 17.0816 | 17 | 0.8961 | 16 | 17 | 17 | 18 | 18 | 1 | -22.1967 | -22 | 0.8706 | -23 | -23 | -22 | -22 | -21 | 1 | 0.033395 | 137.0833 | 1.0833 | 0.036178 | 0 |  | 142.8506 | 123.6902 | 170.5733 | 143 | 124 | 170.5 | 1.193 | 1.1999 | 1.5897 | 141 | 122 | 169 | 142 | 123 | 170 | 143 | 124 | 170 | 143 | 124 | 172 | 144 | 125 | 172 | 1 | 1 | 2
H | 9 | H9 | 1658 | 1082 | 1082 | 1 | 0 | 0 | 133.2874 | 133 | 0.87287 | 132 | 133 | 133 | 134 | 134 | 1 | 34.0055 | 34.5 | 1.6213 | 32 | 32.5 | 34.5 | 35.5 | 35.5 | 3 | 54.5794 | 54.5098 | 0.32707 | 54.1176 | 54.5098 | 54.5098 | 54.902 | 54.902 | 0.39216 | 17.5749 | 18 | 0.78045 | 16 | 17 | 18 | 18 | 18 | 1 | -22.4852 | -23 | 0.92567 | -23 | -23 | -23 | -22 | -21 | 1 | 0.060074 | 135.1231 | 1.1231 | 0.067468 | 0 |  | 141.9464 | 121.6784 | 169.4214 | 142 | 122 | 170 | 1.305 | 1.0967 | 1.8535 | 140 | 120 | 167 | 141 | 121 | 168 | 142 | 122 | 170 | 143 | 122 | 171 | 144 | 123 | 172 | 2 | 1 | 3
H | 10 | H10 | 1664 | 1090 | 1090 | 1 | 0 | 0 | 131.3174 | 131 | 0.84206 | 130 | 131 | 131 | 132 | 132 | 1 | 34.0922 | 34.5 | 1.3732 | 32 | 34 | 34.5 | 34.5 | 35.5 | 0.5 | 53.8482 | 53.7255 | 0.32982 | 53.3333 | 53.7255 | 53.7255 | 54.1176 | 54.1176 | 0.39216 | 17.7982 | 18 | 0.77353 | 17 | 18 | 18 | 18 | 18 | 0 | -21.8165 | -22 | 0.71151 | -23 | -22 | -22 | -21 | -21 | 1 | 0.092661 | 133.1089 | 1.1089 | 0.10275 | 0 |  | 140.9798 | 119.5101 | 166.2248 | 141 | 119 | 166 | 1.212 | 1.0006 | 1.5224 | 139 | 118 | 165 | 140 | 119 | 165 | 141 | 119 | 166 | 142 | 120 | 167 | 142 | 121 | 168 | 2 | 1 | 2
H | 11 | H11 | 1641 | 1068 | 1068 | 1 | 0 | 0 | 128.7322 | 129 | 0.88907 | 128 | 128 | 129 | 129 | 130 | 1 | 33.7477 | 34.5 | 1.7015 | 32 | 34 | 34.5 | 34.5 | 35.5 | 0.5 | 52.8347 | 52.9412 | 0.35233 | 52.549 | 52.549 | 52.9412 | 52.9412 | 53.3333 | 0.39216 | 17.6311 | 18 | 0.95181 | 16 | 18 | 18 | 18 | 18 | 0 | -21.8137 | -22 | 1.143 | -23 | -22 | -22 | -21 | -20 | 1 | 0.027154 | 131.1724 | 1.1724 | 0.031835 | 0 |  | 138.0169 | 117.0599 | 163.5983 | 138 | 117 | 163 | 1.2673 | 1.1537 | 2.2247 | 136 | 116 | 161 | 137 | 116 | 162 | 138 | 117 | 163 | 139 | 118 | 164 | 139 | 119 | 167 | 2 | 2 | 2
H | 12 | H12 | 1606 | 1041 | 1041 | 1 | 0 | 0 | 130.6743 | 130 | 1.389 | 129 | 130 | 130 | 131 | 133 | 1 | 30.9068 | 32 | 1.6856 | 29.5 | 30 | 32 | 32 | 32.5 | 2 | 53.6064 | 53.3333 | 0.56255 | 52.9412 | 53.3333 | 53.3333 | 53.7255 | 54.5098 | 0.39215 | 15.9971 | 16 | 1.011 | 15 | 15 | 16 | 17 | 17 | 2 | -20.4851 | -21 | 0.87487 | -21 | -21 | -21 | -20 | -20 | 1 | 0.03074 | 134.25 | 1.25 | 0.038425 | 0 |  | 138.5841 | 120.1768 | 163.5831 | 139 | 120 | 164 | 1.3448 | 1.7758 | 1.6364 | 137 | 118 | 161 | 138 | 119 | 162 | 139 | 120 | 164 | 139 | 121 | 165 | 140 | 123 | 166 | 1 | 2 | 3

## 05_GEOMETRY_QC
Row | Col | Well | floor_source | local_pitch_px | mouth_r | floor_r | shift_px | shift_frac_of_mouth_r | floor_to_mouth_r_ratio | floor_to_mouth_area_ratio | D_warning | D_critical
A | 1 | A1 | manual_D_projection | 98.2082 | 37.6465 | 37.7723 | 32.6859 | 0.86823 | 1.0033 | 1.0067 | 1 | 1
A | 2 | A2 | manual_D_projection | 97.6033 | 37.4146 | 37.733 | 30.441 | 0.81361 | 1.0085 | 1.0171 | 1 | 1
A | 3 | A3 | manual_D_projection | 97.6033 | 37.4146 | 37.6937 | 28.215 | 0.75412 | 1.0075 | 1.015 | 1 | 1
A | 4 | A4 | manual_D_projection | 97.6033 | 37.4146 | 37.6545 | 26.0537 | 0.69635 | 1.0064 | 1.0129 | 1 | 1
A | 5 | A5 | manual_D_projection | 97.6033 | 37.4146 | 37.6152 | 24.0208 | 0.64202 | 1.0054 | 1.0108 | 1 | 1
A | 6 | A6 | manual_D_projection | 97.6034 | 37.4146 | 37.5759 | 22.2034 | 0.59344 | 1.0043 | 1.0086 | 1 | 1
A | 7 | A7 | manual_D_projection | 97.6034 | 37.4146 | 37.5366 | 20.7159 | 0.55368 | 1.0033 | 1.0065 | 1 | 1
A | 8 | A8 | manual_D_projection | 97.6034 | 37.4146 | 37.4973 | 19.6959 | 0.52642 | 1.0022 | 1.0044 | 1 | 1
A | 9 | A9 | manual_D_projection | 97.6034 | 37.4146 | 37.4581 | 19.2838 | 0.51541 | 1.0012 | 1.0023 | 1 | 1
A | 10 | A10 | manual_D_projection | 97.6033 | 37.4146 | 37.4188 | 19.585 | 0.52346 | 1.0001 | 1.0002 | 1 | 1
A | 11 | A11 | manual_D_projection | 97.6033 | 37.4146 | 37.3795 | 20.6336 | 0.55148 | 0.99906 | 0.99813 | 1 | 1
A | 12 | A12 | manual_D_projection | 97.8055 | 37.4921 | 37.3402 | 22.3865 | 0.5971 | 0.99595 | 0.99192 | 1 | 1
B | 1 | B1 | manual_D_projection | 98.8131 | 37.8784 | 37.7786 | 34.1998 | 0.90288 | 0.99737 | 0.99474 | 1 | 1
B | 2 | B2 | manual_D_projection | 98.1777 | 37.6348 | 37.7393 | 32.236 | 0.85655 | 1.0028 | 1.0056 | 1 | 1
B | 3 | B3 | manual_D_projection | 98.141 | 37.6207 | 37.7 | 30.3086 | 0.80564 | 1.0021 | 1.0042 | 1 | 1
B | 4 | B4 | manual_D_projection | 98.1044 | 37.6067 | 37.6607 | 28.461 | 0.75681 | 1.0014 | 1.0029 | 1 | 1
B | 5 | B5 | manual_D_projection | 98.0678 | 37.5927 | 37.6214 | 26.7493 | 0.71156 | 1.0008 | 1.0015 | 1 | 1
B | 6 | B6 | manual_D_projection | 98.0312 | 37.5786 | 37.5821 | 25.2449 | 0.67179 | 1.0001 | 1.0002 | 1 | 1
B | 7 | B7 | manual_D_projection | 97.9946 | 37.5646 | 37.5428 | 24.0347 | 0.63982 | 0.99942 | 0.99884 | 1 | 1
B | 8 | B8 | manual_D_projection | 97.958 | 37.5506 | 37.5035 | 23.2158 | 0.61826 | 0.99875 | 0.9975 | 1 | 1
B | 9 | B9 | manual_D_projection | 97.9214 | 37.5365 | 37.4642 | 22.8839 | 0.60964 | 0.99807 | 0.99615 | 1 | 1
B | 10 | B10 | manual_D_projection | 97.8848 | 37.5225 | 37.425 | 23.1147 | 0.61602 | 0.9974 | 0.99481 | 1 | 1
B | 11 | B11 | manual_D_projection | 97.8481 | 37.5085 | 37.3857 | 23.9465 | 0.63843 | 0.99673 | 0.99346 | 1 | 1
B | 12 | B12 | manual_D_projection | 98.0077 | 37.5696 | 37.3464 | 25.373 | 0.67536 | 0.99406 | 0.98815 | 1 | 1
C | 1 | C1 | manual_D_projection | 98.8131 | 37.8783 | 37.7848 | 35.8487 | 0.94642 | 0.99753 | 0.99507 | 1 | 1
C | 2 | C2 | manual_D_projection | 98.1837 | 37.6371 | 37.7455 | 34.1356 | 0.90697 | 1.0029 | 1.0058 | 1 | 1
C | 3 | C3 | manual_D_projection | 98.1471 | 37.623 | 37.7062 | 32.467 | 0.86295 | 1.0022 | 1.0044 | 1 | 1
C | 4 | C4 | manual_D_projection | 98.1105 | 37.609 | 37.6669 | 30.8818 | 0.82113 | 1.0015 | 1.0031 | 1 | 1
C | 5 | C5 | manual_D_projection | 98.0739 | 37.595 | 37.6276 | 29.4279 | 0.78276 | 1.0009 | 1.0017 | 1 | 1
C | 6 | C6 | manual_D_projection | 98.0373 | 37.581 | 37.5883 | 28.1634 | 0.74941 | 1.0002 | 1.0004 | 1 | 1
C | 7 | C7 | manual_D_projection | 98.0007 | 37.5669 | 37.549 | 27.1551 | 0.72285 | 0.99952 | 0.99905 | 1 | 1
C | 8 | C8 | manual_D_projection | 97.964 | 37.5529 | 37.5097 | 26.4757 | 0.70502 | 0.99885 | 0.9977 | 1 | 1
C | 9 | C9 | manual_D_projection | 97.9274 | 37.5389 | 37.4704 | 26.1957 | 0.69783 | 0.99818 | 0.99636 | 1 | 1
C | 10 | C10 | manual_D_projection | 97.8908 | 37.5248 | 37.4311 | 26.3742 | 0.70285 | 0.9975 | 0.99501 | 1 | 1
C | 11 | C11 | manual_D_projection | 97.8542 | 37.5108 | 37.3918 | 27.0486 | 0.72109 | 0.99683 | 0.99367 | 1 | 1
C | 12 | C12 | manual_D_projection | 98.0077 | 37.5696 | 37.3525 | 28.2297 | 0.7514 | 0.99422 | 0.98848 | 1 | 1
D | 1 | D1 | manual_D_projection | 98.8131 | 37.8783 | 37.7911 | 37.5349 | 0.99093 | 0.9977 | 0.9954 | 1 | 1
D | 2 | D2 | manual_D_projection | 98.1898 | 37.6394 | 37.7518 | 36.0402 | 0.95751 | 1.003 | 1.006 | 1 | 1
D | 3 | D3 | manual_D_projection | 98.1532 | 37.6254 | 37.7125 | 34.5918 | 0.91937 | 1.0023 | 1.0046 | 1 | 1
D | 4 | D4 | manual_D_projection | 98.1166 | 37.6114 | 37.6732 | 33.2238 | 0.88335 | 1.0016 | 1.0033 | 1 | 1
D | 5 | D5 | manual_D_projection | 98.08 | 37.5973 | 37.6338 | 31.9768 | 0.85051 | 1.001 | 1.0019 | 1 | 1
D | 6 | D6 | manual_D_projection | 98.0434 | 37.5833 | 37.5945 | 30.8982 | 0.82213 | 1.0003 | 1.0006 | 1 | 1
D | 7 | D7 | manual_D_projection | 98.0068 | 37.5693 | 37.5552 | 30.0413 | 0.79963 | 0.99963 | 0.99925 | 1 | 1
D | 8 | D8 | manual_D_projection | 97.9701 | 37.5552 | 37.5159 | 29.4627 | 0.78452 | 0.99895 | 0.99791 | 1 | 1
D | 9 | D9 | manual_D_projection | 97.9335 | 37.5412 | 37.4766 | 29.2177 | 0.77828 | 0.99828 | 0.99656 | 1 | 1
D | 10 | D10 | manual_D_projection | 97.8969 | 37.5272 | 37.4373 | 29.3548 | 0.78223 | 0.99761 | 0.99522 | 1 | 1
D | 11 | D11 | manual_D_projection | 97.8603 | 37.5131 | 37.398 | 29.9095 | 0.79731 | 0.99693 | 0.99387 | 1 | 1
D | 12 | D12 | manual_D_projection | 98.0077 | 37.5696 | 37.3587 | 30.8998 | 0.82247 | 0.99439 | 0.9888 | 1 | 1
E | 1 | E1 | manual_D_projection | 98.8131 | 37.8784 | 37.7974 | 39.1832 | 1.0344 | 0.99786 | 0.99573 | 1 | 1
E | 2 | E2 | manual_D_projection | 98.1959 | 37.6418 | 37.758 | 37.8766 | 1.0062 | 1.0031 | 1.0062 | 1 | 1
E | 3 | E3 | manual_D_projection | 98.1593 | 37.6277 | 37.7187 | 36.6142 | 0.97306 | 1.0024 | 1.0048 | 1 | 1
E | 4 | E4 | manual_D_projection | 98.1227 | 37.6137 | 37.6794 | 35.4258 | 0.94183 | 1.0017 | 1.0035 | 1 | 1
E | 5 | E5 | manual_D_projection | 98.0861 | 37.5997 | 37.6401 | 34.3458 | 0.91346 | 1.0011 | 1.0022 | 1 | 1
E | 6 | E6 | manual_D_projection | 98.0495 | 37.5856 | 37.6007 | 33.4137 | 0.889 | 1.0004 | 1.0008 | 1 | 1
E | 7 | E7 | manual_D_projection | 98.0128 | 37.5716 | 37.5614 | 32.6729 | 0.86962 | 0.99973 | 0.99946 | 1 | 1
E | 8 | E8 | manual_D_projection | 97.9762 | 37.5576 | 37.5221 | 32.1695 | 0.85654 | 0.99906 | 0.99811 | 1 | 1
E | 9 | E9 | manual_D_projection | 97.9397 | 37.5435 | 37.4828 | 31.9487 | 0.85098 | 0.99838 | 0.99677 | 1 | 1
E | 10 | E10 | manual_D_projection | 97.9031 | 37.5295 | 37.4434 | 32.0517 | 0.85404 | 0.99771 | 0.99542 | 1 | 1
E | 11 | E11 | manual_D_projection | 97.8664 | 37.5155 | 37.4041 | 32.5114 | 0.86661 | 0.99703 | 0.99407 | 1 | 1
E | 12 | E12 | manual_D_projection | 98.0076 | 37.5696 | 37.3648 | 33.3495 | 0.88767 | 0.99455 | 0.98913 | 1 | 1
F | 1 | F1 | manual_D_projection | 98.813 | 37.8783 | 37.8036 | 40.7364 | 1.0755 | 0.99803 | 0.99606 | 1 | 1
F | 2 | F2 | manual_D_projection | 98.2021 | 37.6441 | 37.7643 | 39.5908 | 1.0517 | 1.0032 | 1.0064 | 1 | 1
F | 3 | F3 | manual_D_projection | 98.1655 | 37.6301 | 37.725 | 38.4853 | 1.0227 | 1.0025 | 1.005 | 1 | 1
F | 4 | F4 | manual_D_projection | 98.1288 | 37.616 | 37.6856 | 37.4459 | 0.99548 | 1.0018 | 1.0037 | 1 | 1
F | 5 | F5 | manual_D_projection | 98.0922 | 37.602 | 37.6463 | 36.5018 | 0.97074 | 1.0012 | 1.0024 | 1 | 1
F | 6 | F6 | manual_D_projection | 98.0556 | 37.588 | 37.607 | 35.6868 | 0.94942 | 1.0005 | 1.001 | 1 | 1
F | 7 | F7 | manual_D_projection | 98.019 | 37.574 | 37.5676 | 35.0371 | 0.93248 | 0.99983 | 0.99966 | 1 | 1
F | 8 | F8 | manual_D_projection | 97.9824 | 37.5599 | 37.5283 | 34.5911 | 0.92096 | 0.99916 | 0.99832 | 1 | 1
F | 9 | F9 | manual_D_projection | 97.9458 | 37.5459 | 37.4889 | 34.3872 | 0.91587 | 0.99848 | 0.99697 | 1 | 1
F | 10 | F10 | manual_D_projection | 97.9092 | 37.5319 | 37.4496 | 34.4612 | 0.91818 | 0.99781 | 0.99562 | 1 | 1
F | 11 | F11 | manual_D_projection | 97.8726 | 37.5178 | 37.4103 | 34.8435 | 0.92872 | 0.99713 | 0.99427 | 1 | 1
F | 12 | F12 | manual_D_projection | 98.0076 | 37.5696 | 37.3709 | 35.5571 | 0.94643 | 0.99471 | 0.98945 | 1 | 1
G | 1 | G1 | manual_D_projection | 98.813 | 37.8783 | 37.8099 | 42.1511 | 1.1128 | 0.99819 | 0.99639 | 1 | 1
G | 2 | G2 | manual_D_projection | 98.2083 | 37.6465 | 37.7705 | 41.1432 | 1.0929 | 1.0033 | 1.0066 | 1 | 1
G | 3 | G3 | manual_D_projection | 98.1716 | 37.6325 | 37.7312 | 40.1704 | 1.0674 | 1.0026 | 1.0053 | 1 | 1
G | 4 | G4 | manual_D_projection | 98.135 | 37.6184 | 37.6919 | 39.255 | 1.0435 | 1.002 | 1.0039 | 1 | 1
G | 5 | G5 | manual_D_projection | 98.0984 | 37.6044 | 37.6525 | 38.423 | 1.0218 | 1.0013 | 1.0026 | 1 | 1
G | 6 | G6 | manual_D_projection | 98.0618 | 37.5904 | 37.6132 | 37.7027 | 1.003 | 1.0006 | 1.0012 | 1 | 1
G | 7 | G7 | manual_D_projection | 98.0252 | 37.5763 | 37.5738 | 37.1257 | 0.98801 | 0.99993 | 0.99987 | 1 | 1
G | 8 | G8 | manual_D_projection | 97.9886 | 37.5623 | 37.5345 | 36.7245 | 0.9777 | 0.99926 | 0.99852 | 1 | 1
G | 9 | G9 | manual_D_projection | 97.952 | 37.5483 | 37.4951 | 36.5322 | 0.97294 | 0.99858 | 0.99717 | 1 | 1
G | 10 | G10 | manual_D_projection | 97.9154 | 37.5342 | 37.4558 | 36.5807 | 0.9746 | 0.99791 | 0.99582 | 1 | 1
G | 11 | G11 | manual_D_projection | 97.8787 | 37.5202 | 37.4164 | 36.8984 | 0.98343 | 0.99723 | 0.99448 | 1 | 1
G | 12 | G12 | manual_D_projection | 98.0077 | 37.5696 | 37.3771 | 37.5082 | 0.99837 | 0.99488 | 0.98978 | 1 | 1
H | 1 | H1 | manual_D_projection | 98.2511 | 37.6629 | 37.8162 | 43.3945 | 1.1522 | 1.0041 | 1.0082 | 1 | 1
H | 2 | H2 | manual_D_projection | 97.6891 | 37.4475 | 37.7768 | 42.5044 | 1.135 | 1.0088 | 1.0177 | 1 | 1
H | 3 | H3 | manual_D_projection | 97.6891 | 37.4475 | 37.7374 | 41.6438 | 1.1121 | 1.0077 | 1.0155 | 1 | 1
H | 4 | H4 | manual_D_projection | 97.6891 | 37.4475 | 37.6981 | 40.8327 | 1.0904 | 1.0067 | 1.0134 | 1 | 1
H | 5 | H5 | manual_D_projection | 97.6891 | 37.4475 | 37.6587 | 40.0934 | 1.0707 | 1.0056 | 1.0113 | 1 | 1
H | 6 | H6 | manual_D_projection | 97.6891 | 37.4475 | 37.6194 | 39.4511 | 1.0535 | 1.0046 | 1.0092 | 1 | 1
H | 7 | H7 | manual_D_projection | 97.6891 | 37.4475 | 37.58 | 38.9326 | 1.0397 | 1.0035 | 1.0071 | 1 | 1
H | 8 | H8 | manual_D_projection | 97.6891 | 37.4475 | 37.5407 | 38.5666 | 1.0299 | 1.0025 | 1.005 | 1 | 1
H | 9 | H9 | manual_D_projection | 97.6891 | 37.4475 | 37.5013 | 38.3823 | 1.025 | 1.0014 | 1.0029 | 1 | 1
H | 10 | H10 | manual_D_projection | 97.6891 | 37.4475 | 37.4619 | 38.4082 | 1.0257 | 1.0004 | 1.0008 | 1 | 1
H | 11 | H11 | manual_D_projection | 97.6891 | 37.4475 | 37.4226 | 38.6708 | 1.0327 | 0.99934 | 0.99867 | 1 | 1
H | 12 | H12 | manual_D_projection | 97.8484 | 37.5085 | 37.3832 | 39.1926 | 1.0449 | 0.99666 | 0.99333 | 1 | 1

## 06_WELL_BOTTOM
Row | Col | Well | cx | cy | local_pitch_px | px_per_mm | cyl_r_bg | mouth_r_geom | floor_r_geom | mouth_cx | mouth_cy | mouth_r | mouth_score | floor_cx | floor_cy | floor_r | floor_score | shift_px
A | 1 | A1 | 185.8214 | 143.5893 | 98.2082 | 10.912 | 41.3969 | 37.6465 | 35.355 | 185.8214 | 143.5893 | 37.6465 | 0.47652 | 211.1607 | 122.9425 | 37.7723 |  | 32.6859
A | 2 | A2 | 283.4246 | 143.7599 | 97.6033 | 10.8448 | 41.3969 | 37.4146 | 35.1372 | 283.4246 | 143.7599 | 37.4146 | 8.4875 | 305.9517 | 123.2859 | 37.733 |  | 30.441
A | 3 | A3 | 381.0278 | 143.9306 | 97.6033 | 10.8448 | 41.3969 | 37.4146 | 35.1372 | 381.0278 | 143.9306 | 37.4146 | 10.7546 | 400.6219 | 123.6289 | 37.6937 |  | 28.215
A | 4 | A4 | 478.631 | 144.1012 | 97.6033 | 10.8448 | 41.3969 | 37.4146 | 35.1372 | 478.631 | 144.1012 | 37.4146 | 12.4513 | 495.1715 | 123.9715 | 37.6545 |  | 26.0537
A | 5 | A5 | 576.2341 | 144.2718 | 97.6033 | 10.8448 | 41.5907 | 37.4146 | 35.1372 | 576.2341 | 144.2718 | 37.4146 | 14.5922 | 589.6008 | 124.3136 | 37.6152 |  | 24.0208
A | 6 | A6 | 673.8373 | 144.4425 | 97.6034 | 10.8448 | 42 | 37.4146 | 35.1372 | 673.8373 | 144.4425 | 37.4146 | 15.6902 | 683.91 | 124.6553 | 37.5759 |  | 22.2034
A | 7 | A7 | 771.4405 | 144.6131 | 97.6034 | 10.8448 | 41.3969 | 37.4146 | 35.1372 | 771.4405 | 144.6131 | 37.4146 | 19.3702 | 778.0993 | 124.9966 | 37.5366 |  | 20.7159
A | 8 | A8 | 869.0436 | 144.7837 | 97.6034 | 10.8448 | 41.3969 | 37.4146 | 35.1372 | 869.0436 | 144.7837 | 37.4146 | 19.7131 | 872.169 | 125.3374 | 37.4973 |  | 19.6959
A | 9 | A9 | 966.6469 | 144.9544 | 97.6034 | 10.8448 | 41.3752 | 37.4146 | 35.1372 | 966.6469 | 144.9544 | 37.4146 | 20.1582 | 966.1193 | 125.6778 | 37.4581 |  | 19.2838
A | 10 | A10 | 1064.25 | 145.125 | 97.6033 | 10.8448 | 41.3969 | 37.4146 | 35.1372 | 1064.25 | 145.125 | 37.4146 | 23.1261 | 1059.9504 | 126.0178 | 37.4188 |  | 19.585
A | 11 | A11 | 1161.8531 | 145.2956 | 97.6033 | 10.8448 | 41.3969 | 37.4146 | 35.1372 | 1161.8531 | 145.2956 | 37.4146 | 22.5024 | 1153.6626 | 126.3573 | 37.3795 |  | 20.6336
A | 12 | A12 | 1259.4563 | 145.4663 | 97.8055 | 10.8673 | 44.369 | 37.4921 | 35.21 | 1259.4563 | 145.4663 | 37.4921 | 15.8891 | 1247.256 | 126.6964 | 37.3402 |  | 22.3865
B | 1 | B1 | 185.017 | 242.3991 | 98.8131 | 10.9792 | 42 | 37.8784 | 35.5727 | 185.017 | 242.3991 | 37.8784 | 0.59996 | 209.0354 | 218.0528 | 37.7786 |  | 34.1998
B | 2 | B2 | 282.6324 | 242.4966 | 98.1777 | 10.9086 | 41.5907 | 37.6348 | 35.344 | 282.6324 | 242.4966 | 37.6348 | 8.3259 | 303.9731 | 218.336 | 37.7393 |  | 32.236
B | 3 | B3 | 380.2477 | 242.5941 | 98.141 | 10.9046 | 41.3969 | 37.6207 | 35.3308 | 380.2477 | 242.5941 | 37.6207 | 10.2069 | 398.7896 | 218.6189 | 37.7 |  | 30.3086
B | 4 | B4 | 477.8631 | 242.6916 | 98.1044 | 10.9005 | 41.3969 | 37.6067 | 35.3176 | 477.8631 | 242.6916 | 37.6067 | 11.8937 | 493.4853 | 218.9014 | 37.6607 |  | 28.461
B | 5 | B5 | 575.4784 | 242.7891 | 98.0678 | 10.8964 | 41.1814 | 37.5927 | 35.3044 | 575.4784 | 242.7891 | 37.5927 | 13.7705 | 588.0602 | 219.1836 | 37.6214 |  | 26.7493
B | 6 | B6 | 673.0938 | 242.8866 | 98.0312 | 10.8924 | 41.3969 | 37.5786 | 35.2912 | 673.0938 | 242.8866 | 37.5786 | 15.1962 | 682.5146 | 219.4654 | 37.5821 |  | 25.2449
B | 7 | B7 | 770.7092 | 242.9841 | 97.9946 | 10.8883 | 41.9783 | 37.5646 | 35.278 | 770.7092 | 242.9841 | 37.5646 | 18.6168 | 776.8489 | 219.7468 | 37.5428 |  | 24.0347
B | 8 | B8 | 868.3245 | 243.0816 | 97.958 | 10.8842 | 41.3969 | 37.5506 | 35.2649 | 868.3245 | 243.0816 | 37.5506 | 20.1675 | 871.0632 | 220.0279 | 37.5035 |  | 23.2158
B | 9 | B9 | 965.9398 | 243.1791 | 97.9214 | 10.8802 | 41.3969 | 37.5365 | 35.2517 | 965.9398 | 243.1791 | 37.5365 | 21.8732 | 965.1576 | 220.3086 | 37.4642 |  | 22.8839
B | 10 | B10 | 1063.5552 | 243.2766 | 97.8848 | 10.8761 | 41.3969 | 37.5225 | 35.2385 | 1063.5552 | 243.2766 | 37.5225 | 22.7684 | 1059.1326 | 220.589 | 37.425 |  | 23.1147
B | 11 | B11 | 1161.1705 | 243.3742 | 97.8481 | 10.872 | 41.9783 | 37.5085 | 35.2253 | 1161.1705 | 243.3742 | 37.5085 | 22.2147 | 1152.9882 | 220.869 | 37.3857 |  | 23.9465
B | 12 | B12 | 1258.7859 | 243.4716 | 98.0077 | 10.8897 | 45.1659 | 37.5696 | 35.2828 | 1258.7859 | 243.4716 | 37.5696 | 16.2566 | 1246.7246 | 221.1487 | 37.3464 |  | 25.373
C | 1 | C1 | 184.2126 | 341.2089 | 98.8131 | 10.9792 | 51.5411 | 37.8783 | 35.5727 | 184.2126 | 341.2089 | 37.8783 | 1.2359 | 206.9034 | 313.4555 | 37.7848 |  | 35.8487
C | 2 | C2 | 281.8401 | 341.2333 | 98.1837 | 10.9093 | 46.9535 | 37.6371 | 35.3461 | 281.8401 | 341.2333 | 37.6371 | 8.499 | 301.9884 | 313.6782 | 37.7455 |  | 34.1356
C | 3 | C3 | 379.4677 | 341.2577 | 98.1471 | 10.9052 | 41.3969 | 37.623 | 35.3329 | 379.4677 | 341.2577 | 37.623 | 10.784 | 396.9518 | 313.9005 | 37.7062 |  | 32.467
C | 4 | C4 | 477.0952 | 341.282 | 98.1105 | 10.9012 | 41.9783 | 37.609 | 35.3198 | 477.0952 | 341.282 | 37.609 | 12.5083 | 491.7939 | 314.1227 | 37.6669 |  | 30.8818
C | 5 | C5 | 574.7228 | 341.3064 | 98.0739 | 10.8971 | 41.3969 | 37.595 | 35.3066 | 574.7228 | 341.3064 | 37.595 | 13.9024 | 586.5149 | 314.3445 | 37.6276 |  | 29.4279
C | 6 | C6 | 672.3503 | 341.3308 | 98.0373 | 10.893 | 41.3969 | 37.581 | 35.2934 | 672.3503 | 341.3308 | 37.581 | 14.6913 | 681.1151 | 314.566 | 37.5883 |  | 28.1634
C | 7 | C7 | 769.9778 | 341.3552 | 98.0007 | 10.889 | 41.3969 | 37.5669 | 35.2802 | 769.9778 | 341.3552 | 37.5669 | 17.76 | 775.5947 | 314.7873 | 37.549 |  | 27.1551
C | 8 | C8 | 867.6054 | 341.3795 | 97.964 | 10.8849 | 41.3969 | 37.5529 | 35.2671 | 867.6054 | 341.3795 | 37.5529 | 20.0774 | 869.9539 | 315.0082 | 37.5097 |  | 26.4757
C | 9 | C9 | 965.233 | 341.4039 | 97.9274 | 10.8808 | 41.9783 | 37.5389 | 35.2539 | 965.233 | 341.4039 | 37.5389 | 20.9368 | 964.1929 | 315.2289 | 37.4704 |  | 26.1957
C | 10 | C10 | 1062.8605 | 341.4283 | 97.8908 | 10.8768 | 41.3969 | 37.5248 | 35.2407 | 1062.8605 | 341.4283 | 37.5248 | 21.8795 | 1058.3121 | 315.4493 | 37.4311 |  | 26.3742
C | 11 | C11 | 1160.488 | 341.4527 | 97.8542 | 10.8727 | 41.9783 | 37.5108 | 35.2275 | 1160.488 | 341.4527 | 37.5108 | 22.0922 | 1152.3115 | 315.6695 | 37.3918 |  | 27.0486
C | 12 | C12 | 1258.1156 | 341.4771 | 98.0077 | 10.8897 | 45.1659 | 37.5696 | 35.2828 | 1258.1156 | 341.4771 | 37.5696 | 14.6045 | 1246.1915 | 315.8893 | 37.3525 |  | 28.2297
D | 1 | D1 | 183.4082 | 440.0187 | 98.8131 | 10.9792 | 42 | 37.8783 | 35.5727 | 183.4082 | 440.0187 | 37.8783 | 1.2276 | 204.765 | 409.1519 | 37.7911 |  | 37.5349
D | 2 | D2 | 281.0479 | 439.97 | 98.1898 | 10.91 | 56.9256 | 37.6394 | 35.3483 | 281.0479 | 439.97 | 37.6394 | 8.5701 | 299.9976 | 409.3137 | 37.7518 |  | 36.0402
D | 3 | D3 | 378.6877 | 439.9212 | 98.1532 | 10.9059 | 42 | 37.6254 | 35.3352 | 378.6877 | 439.9212 | 37.6254 | 10.6954 | 395.1082 | 409.4752 | 37.7125 |  | 34.5918
D | 4 | D4 | 476.3274 | 439.8725 | 98.1166 | 10.9018 | 41.3969 | 37.6114 | 35.322 | 476.3274 | 439.8725 | 37.6114 | 12.7021 | 490.0973 | 409.6365 | 37.6732 |  | 33.2238
D | 5 | D5 | 573.9672 | 439.8237 | 98.08 | 10.8978 | 41.3969 | 37.5973 | 35.3088 | 573.9672 | 439.8237 | 37.5973 | 15.7217 | 584.9648 | 409.7976 | 37.6338 |  | 31.9768
D | 6 | D6 | 671.6069 | 439.775 | 98.0434 | 10.8937 | 41.3969 | 37.5833 | 35.2956 | 671.6069 | 439.775 | 37.5833 | 16.6462 | 679.7112 | 409.9585 | 37.5945 |  | 30.8982
D | 7 | D7 | 769.2466 | 439.7262 | 98.0068 | 10.8896 | 41.3969 | 37.5693 | 35.2824 | 769.2466 | 439.7262 | 37.5693 | 17.7508 | 774.3366 | 410.1192 | 37.5552 |  | 30.0413
D | 8 | D8 | 866.8864 | 439.6774 | 97.9701 | 10.8856 | 41.5907 | 37.5552 | 35.2693 | 866.8864 | 439.6774 | 37.5552 | 20.769 | 868.8412 | 410.2797 | 37.5159 |  | 29.4627
D | 9 | D9 | 964.5261 | 439.6287 | 97.9335 | 10.8815 | 40.6 | 37.5412 | 35.2561 | 964.5261 | 439.6287 | 37.5412 | 20.6074 | 963.2253 | 410.44 | 37.4766 |  | 29.2177
D | 10 | D10 | 1062.1658 | 439.58 | 97.8969 | 10.8774 | 41.3752 | 37.5272 | 35.2429 | 1062.1658 | 439.58 | 37.5272 | 22.1939 | 1057.4891 | 410.6 | 37.4373 |  | 29.3548
D | 11 | D11 | 1159.8055 | 439.5312 | 97.8603 | 10.8734 | 40.6 | 37.5131 | 35.2297 | 1159.8055 | 439.5312 | 37.5131 | 24.5828 | 1151.6329 | 410.7599 | 37.398 |  | 29.9095
D | 12 | D12 | 1257.4453 | 439.4824 | 98.0077 | 10.8897 | 49.1504 | 37.5696 | 35.2828 | 1257.4453 | 439.4824 | 37.5696 | 15.4755 | 1245.657 | 410.9196 | 37.3587 |  | 30.8998
E | 1 | E1 | 182.6037 | 538.8286 | 98.8131 | 10.9792 | 41.3969 | 37.8784 | 35.5727 | 182.6037 | 538.8286 | 37.8784 | 1.0372 | 202.6199 | 505.1436 | 37.7974 |  | 39.1832
E | 2 | E2 | 280.2557 | 538.7067 | 98.1959 | 10.9107 | 60.1194 | 37.6418 | 35.3505 | 280.2557 | 538.7067 | 37.6418 | 13.3896 | 298.0006 | 505.244 | 37.758 |  | 37.8766
E | 3 | E3 | 377.9076 | 538.5848 | 98.1593 | 10.9066 | 41.3969 | 37.6277 | 35.3374 | 377.9076 | 538.5848 | 37.6277 | 14.4426 | 393.2591 | 505.3442 | 37.7187 |  | 36.6142
E | 4 | E4 | 475.5595 | 538.463 | 98.1227 | 10.9025 | 42 | 37.6137 | 35.3242 | 475.5595 | 538.463 | 37.6137 | 15.2901 | 488.3954 | 505.4444 | 37.6794 |  | 35.4258
E | 5 | E5 | 573.2114 | 538.3411 | 98.0861 | 10.8985 | 41.9783 | 37.5997 | 35.311 | 573.2114 | 538.3411 | 37.5997 | 15.9695 | 583.41 | 505.5444 | 37.6401 |  | 34.3458
E | 6 | E6 | 670.8634 | 538.2192 | 98.0495 | 10.8944 | 42 | 37.5856 | 35.2978 | 670.8634 | 538.2192 | 37.5856 | 19.2487 | 678.303 | 505.6443 | 37.6007 |  | 33.4137
E | 7 | E7 | 768.5153 | 538.0973 | 98.0128 | 10.8903 | 42.7752 | 37.5716 | 35.2846 | 768.5153 | 538.0973 | 37.5716 | 19.9587 | 773.0746 | 505.744 | 37.5614 |  | 32.6729
E | 8 | E8 | 866.1672 | 537.9754 | 97.9762 | 10.8862 | 41.7845 | 37.5576 | 35.2714 | 866.1672 | 537.9754 | 37.5576 | 21.0179 | 867.7252 | 505.8437 | 37.5221 |  | 32.1695
E | 9 | E9 | 963.8192 | 537.8535 | 97.9397 | 10.8822 | 41.3969 | 37.5435 | 35.2583 | 963.8192 | 537.8535 | 37.5435 | 22.2575 | 962.2548 | 505.9431 | 37.4828 |  | 31.9487
E | 10 | E10 | 1061.4711 | 537.7316 | 97.9031 | 10.8781 | 41.3969 | 37.5295 | 35.2451 | 1061.4711 | 537.7316 | 37.5295 | 21.865 | 1056.6637 | 506.0425 | 37.4434 |  | 32.0517
E | 11 | E11 | 1159.123 | 537.6097 | 97.8664 | 10.874 | 41.3752 | 37.5155 | 35.2319 | 1159.123 | 537.6097 | 37.5155 | 21.886 | 1150.9523 | 506.1418 | 37.4041 |  | 32.5114
E | 12 | E12 | 1256.7749 | 537.4879 | 98.0076 | 10.8897 | 54.7287 | 37.5696 | 35.2827 | 1256.7749 | 537.4879 | 37.5696 | 16.0546 | 1245.1207 | 506.2409 | 37.3648 |  | 33.3495
F | 1 | F1 | 181.7993 | 637.6384 | 98.813 | 10.9792 | 41.7845 | 37.8783 | 35.5727 | 181.7993 | 637.6384 | 37.8783 | 0.73362 | 200.4682 | 601.4316 | 37.8036 |  | 40.7364
F | 2 | F2 | 279.4634 | 637.4434 | 98.2021 | 10.9113 | 62.1194 | 37.6441 | 35.3527 | 279.4634 | 637.4434 | 37.6441 | 12.5395 | 295.9975 | 601.4703 | 37.7643 |  | 39.5908
F | 3 | F3 | 377.1276 | 637.2483 | 98.1655 | 10.9073 | 41.9783 | 37.6301 | 35.3396 | 377.1276 | 637.2483 | 37.6301 | 13.3593 | 391.4042 | 601.509 | 37.725 |  | 38.4853
F | 4 | F4 | 474.7917 | 637.0533 | 98.1288 | 10.9032 | 42 | 37.616 | 35.3264 | 474.7917 | 637.0533 | 37.616 | 14.7703 | 486.6884 | 601.5475 | 37.6856 |  | 37.4459
F | 5 | F5 | 572.4558 | 636.8583 | 98.0922 | 10.8991 | 41.3969 | 37.602 | 35.3132 | 572.4558 | 636.8583 | 37.602 | 16.6174 | 581.8504 | 601.5861 | 37.6463 |  | 36.5018
F | 6 | F6 | 670.1199 | 636.6633 | 98.0556 | 10.8951 | 41.5907 | 37.588 | 35.3 | 670.1199 | 636.6633 | 37.588 | 18.1677 | 676.8904 | 601.6246 | 37.607 |  | 35.6868
F | 7 | F7 | 767.7841 | 636.4683 | 98.019 | 10.891 | 41.3969 | 37.574 | 35.2868 | 767.7841 | 636.4683 | 37.574 | 19.763 | 771.8088 | 601.6631 | 37.5676 |  | 35.0371
F | 8 | F8 | 865.4481 | 636.2732 | 97.9824 | 10.8869 | 43.4 | 37.5599 | 35.2737 | 865.4481 | 636.2732 | 37.5599 | 22.098 | 866.6057 | 601.7015 | 37.5283 |  | 34.5911
F | 9 | F9 | 963.1123 | 636.0782 | 97.9458 | 10.8829 | 41.7845 | 37.5459 | 35.2605 | 963.1123 | 636.0782 | 37.5459 | 20.7295 | 961.2812 | 601.7398 | 37.4889 |  | 34.3872
F | 10 | F10 | 1060.7764 | 635.8832 | 97.9092 | 10.8788 | 41.3969 | 37.5319 | 35.2473 | 1060.7764 | 635.8832 | 37.5319 | 21.6635 | 1055.8358 | 601.7781 | 37.4496 |  | 34.4612
F | 11 | F11 | 1158.4406 | 635.6882 | 97.8726 | 10.8747 | 41.3969 | 37.5178 | 35.2341 | 1158.4406 | 635.6882 | 37.5178 | 20.4248 | 1150.2695 | 601.8163 | 37.4103 |  | 34.8435
F | 12 | F12 | 1256.1046 | 635.4932 | 98.0076 | 10.8897 | 56.3225 | 37.5696 | 35.2827 | 1256.1046 | 635.4932 | 37.5696 | 16.7596 | 1244.5828 | 601.8546 | 37.3709 |  | 35.5571
G | 1 | G1 | 180.9949 | 736.4481 | 98.813 | 10.9792 | 41.9783 | 37.8783 | 35.5727 | 180.9949 | 736.4481 | 37.8783 | 0.91739 | 198.3099 | 698.0176 | 37.8099 |  | 42.1511
G | 2 | G2 | 278.6712 | 736.18 | 98.2083 | 10.912 | 62.5132 | 37.6465 | 35.355 | 278.6712 | 736.18 | 37.6465 | 12.552 | 293.9882 | 697.9942 | 37.7705 |  | 41.1432
G | 3 | G3 | 376.3475 | 735.9119 | 98.1716 | 10.908 | 41.3969 | 37.6325 | 35.3418 | 376.3475 | 735.9119 | 37.6325 | 13.4452 | 389.5435 | 697.9708 | 37.7312 |  | 40.1704
G | 4 | G4 | 474.0239 | 735.6437 | 98.135 | 10.9039 | 41.3969 | 37.6184 | 35.3286 | 474.0239 | 735.6437 | 37.6184 | 14.4937 | 484.976 | 697.9475 | 37.6919 |  | 39.255
G | 5 | G5 | 571.7001 | 735.3756 | 98.0984 | 10.8998 | 42 | 37.6044 | 35.3154 | 571.7001 | 735.3756 | 37.6044 | 15.3979 | 580.2859 | 697.9242 | 37.6525 |  | 38.423
G | 6 | G6 | 669.3765 | 735.1074 | 98.0618 | 10.8958 | 41.5907 | 37.5904 | 35.3022 | 669.3765 | 735.1074 | 37.5904 | 16.699 | 675.4736 | 697.9009 | 37.6132 |  | 37.7027
G | 7 | G7 | 767.0528 | 734.8393 | 98.0252 | 10.8917 | 41.7845 | 37.5763 | 35.2891 | 767.0528 | 734.8393 | 37.5763 | 19.2174 | 770.5391 | 697.8777 | 37.5738 |  | 37.1257
G | 8 | G8 | 864.7291 | 734.5712 | 97.9886 | 10.8876 | 46.5659 | 37.5623 | 35.2759 | 864.7291 | 734.5712 | 37.5623 | 21.1226 | 865.4827 | 697.8544 | 37.5345 |  | 36.7245
G | 9 | G9 | 962.4054 | 734.303 | 97.952 | 10.8836 | 42 | 37.5483 | 35.2627 | 962.4054 | 734.303 | 37.5483 | 19.1542 | 960.3047 | 697.8313 | 37.4951 |  | 36.5322
G | 10 | G10 | 1060.0818 | 734.0349 | 97.9154 | 10.8795 | 41.3969 | 37.5342 | 35.2496 | 1060.0818 | 734.0349 | 37.5342 | 18.5947 | 1055.0052 | 697.8081 | 37.4558 |  | 36.5807
G | 11 | G11 | 1157.7581 | 733.7668 | 97.8787 | 10.8754 | 42.1938 | 37.5202 | 35.2363 | 1157.7581 | 733.7668 | 37.5202 | 19.4934 | 1149.5847 | 697.785 | 37.4164 |  | 36.8984
G | 12 | G12 | 1255.4343 | 733.4986 | 98.0077 | 10.8897 | 59.3163 | 37.5696 | 35.2828 | 1255.4343 | 733.4986 | 37.5696 | 16.561 | 1244.0432 | 697.7619 | 37.3771 |  | 37.5082
H | 1 | H1 | 180.1905 | 835.2579 | 98.2511 | 10.9168 | 41.9783 | 37.6629 | 35.3704 | 180.1905 | 835.2579 | 37.6629 | 0.92849 | 196.1448 | 794.9028 | 37.8162 |  | 43.3945
H | 2 | H2 | 277.879 | 834.9167 | 97.6891 | 10.8543 | 66.1194 | 37.4475 | 35.1681 | 277.879 | 834.9167 | 37.4475 | 8.3003 | 291.9727 | 794.8169 | 37.7768 |  | 42.5044
H | 3 | H3 | 375.5675 | 834.5754 | 97.6891 | 10.8543 | 62.1163 | 37.4475 | 35.1681 | 375.5675 | 834.5754 | 37.4475 | 8.6381 | 387.6772 | 794.7311 | 37.7374 |  | 41.6438
H | 4 | H4 | 473.256 | 834.2341 | 97.6891 | 10.8543 | 63.3225 | 37.4475 | 35.1681 | 473.256 | 834.2341 | 37.4475 | 9.7012 | 483.2584 | 794.6455 | 37.6981 |  | 40.8327
H | 5 | H5 | 570.9445 | 833.8928 | 97.6891 | 10.8543 | 65.7349 | 37.4475 | 35.1681 | 570.9445 | 833.8928 | 37.4475 | 9.6762 | 578.7167 | 794.56 | 37.6587 |  | 40.0934
H | 6 | H6 | 668.6329 | 833.5516 | 97.6891 | 10.8543 | 72.1101 | 37.4475 | 35.1681 | 668.6329 | 833.5516 | 37.4475 | 10.7356 | 674.0523 | 794.4745 | 37.6194 |  | 39.4511
H | 7 | H7 | 766.3215 | 833.2103 | 97.6891 | 10.8543 | 67.3721 | 37.4475 | 35.1681 | 766.3215 | 833.2103 | 37.4475 | 10.5482 | 769.2654 | 794.3892 | 37.58 |  | 38.9326
H | 8 | H8 | 864.0099 | 832.869 | 97.6891 | 10.8543 | 65.1752 | 37.4475 | 35.1681 | 864.0099 | 832.869 | 37.4475 | 11.1296 | 864.3563 | 794.304 | 37.5407 |  | 38.5666
H | 9 | H9 | 961.6985 | 832.5278 | 97.6891 | 10.8543 | 64.7659 | 37.4475 | 35.1681 | 961.6985 | 832.5278 | 37.4475 | 11.6475 | 959.3251 | 794.2189 | 37.5013 |  | 38.3823
H | 10 | H10 | 1059.387 | 832.1865 | 97.6891 | 10.8543 | 67.5442 | 37.4475 | 35.1681 | 1059.387 | 832.1865 | 37.4475 | 12.3155 | 1054.1722 | 794.134 | 37.4619 |  | 38.4082
H | 11 | H11 | 1157.0754 | 831.8453 | 97.6891 | 10.8543 | 62.138 | 37.4475 | 35.1681 | 1157.0754 | 831.8453 | 37.4475 | 12.2742 | 1148.8977 | 794.0491 | 37.4226 |  | 38.6708
H | 12 | H12 | 1254.7639 | 831.504 | 97.8484 | 10.872 | 83.569 | 37.5085 | 35.2254 | 1254.7639 | 831.504 | 37.5085 | 11.7447 | 1243.502 | 793.9643 | 37.3832 |  | 39.1926

## 07_PLATE_GEOMETRY
key | value
a1_col_offset_mm | 14.38
a1_col_offset_mm_range | (14.29, 14.38)
a1_row_offset_mm | 11.24
a1_row_offset_mm_range | (11.18, 11.24)
bridge_width_mm | 0.6
extra_optical_margin_mm | 0.2
flange_or_skirt_height_mm | 2.5
floor_diam_mm | 6.48
floor_diam_mm_range | (6.35, 6.58)
footprint_length_mm | 127.76
footprint_length_mm_range | (127.76, 127.8)
footprint_width_mm | 85.48
footprint_width_mm_range | (85.48, 85.5)
inner_diam_mm | 6.9
mouth_diam_mm | 6.9
mouth_diam_mm_range | (6.69, 6.96)
name | 96-well standard F-bottom / flat-bottom
outer_diam_mm | 7.75
pitch_mm | 9
pitch_mm_range | (9.0, 9.02)
plate_height_mm_nominal | 14.4
plate_height_mm_range | (14.2, 14.6)
standard_basis | ANSI/SLAS-compatible flat-bottom geometry; vendor dimensions may vary by plate family
well_bottom_elevation_mm_nominal | 3.6
well_bottom_elevation_mm_range | (3.5, 3.7)
well_depth_mm | 10.9
well_depth_mm_range | (10.67, 10.9)

## 08_EMPTY_WELLS
Row | Col | Well | MeanW_Red | MeanBG_Red | PAbs_Red | MeanW_Green | MeanBG_Green | PAbs_Green | MeanW_Blue | MeanBG_Blue | PAbs_Blue | L | a | b | UsedFraction | BrightExcludedFraction | HighlightIndex
A | 1 | A1 | 0.64556 | 0.63923 | -0.0042734 | 0.65922 | 0.6483 | -0.0072593 | 0.63204 | 0.652 | 0.013501 | 84.3137 | -1 | 1 | 1 | 0.036755 | 0.036755
B | 1 | B1 | 0.64556 | 0.63529 | -0.0069585 | 0.65922 | 0.63818 | -0.014089 | 0.63878 | 0.64674 | 0.0053797 | 84.3137 | -1 | 1 | 1 | 0.013858 | 0.016776
C | 1 | C1 | 0.64556 | 0.63132 | -0.0096851 | 0.65922 | 0.62978 | -0.019843 | 0.63878 | 0.63997 | 0.00081057 | 84.3137 | -1 | 1 | 1 | 0.010172 | 0.010172
D | 1 | D1 | 0.63878 | 0.62731 | -0.0078708 | 0.65237 | 0.62306 | -0.019962 | 0.63204 | 0.63172 | -0.00022129 | 83.9216 | -1 | 1 | 1 | 0.074766 | 0.076466
E | 1 | E1 | 0.63878 | 0.62326 | -0.010681 | 0.64556 | 0.618 | -0.018946 | 0.62534 | 0.62202 | -0.0023171 | 83.5294 | -1 | 1 | 1 | 0.012832 | 0.012832
F | 1 | F1 | 0.65237 | 0.61918 | -0.022677 | 0.65237 | 0.61456 | -0.025926 | 0.62534 | 0.6109 | -0.010149 | 84.3137 | -1 | 2 | 1 | 0.043435 | 0.04541
G | 1 | G1 | 0.63878 | 0.61507 | -0.016429 | 0.63878 | 0.61275 | -0.01807 | 0.61207 | 0.59842 | -0.0097951 | 83.5294 | -1 | 2 | 1 | 0.026681 | 0.032017
H | 1 | H1 | 0.63878 | 0.61092 | -0.019368 | 0.63878 | 0.61254 | -0.018218 | 0.61207 | 0.58461 | -0.019931 | 83.5294 | -1 | 2 | 1 | 0.071926 | 0.071926

## 09_SPATIAL_DIAGNOSTICS
Dataset | Status | Applicability | Reason | n | intercept | slope_col | slope_row | R2 | corr_col | corr_row
unknown | not_applied | requires usable unknown or empty wells distributed across rows/columns | no usable wells available for this spatial-trend dataset | 0
empty | not_applied | requires usable unknown or empty wells distributed across rows/columns | no usable wells available for this spatial-trend dataset | 0

## 10_METHOD_COMPARISON
Method | Family | ComparableGroup | CommonFactorsN | Score | ScoreFormula | RankMode | R2_cal | R2_std_mean | m_cal | m_std_mean | SlopeAgreement | beta_mean | bias_index_mean | C0_median | C0_sd_median | Estimate_value | Estimate_sd | Estimate_source | LOD | LOQ | SNR | n_stdadd | n_unknown | C0_mean | expected_label_Colorimetry_4 | expected_id_Colorimetry_4 | expected_value_Colorimetry_4 | expected_sd_Colorimetry_4 | estimate_for_expected_Colorimetry_4 | delta_expected_Colorimetry_4 | recovery_pct_Colorimetry_4 | rel_error_Colorimetry_4 | expected_label_Colorimetry_8 | expected_id_Colorimetry_8 | expected_value_Colorimetry_8 | expected_sd_Colorimetry_8 | estimate_for_expected_Colorimetry_8 | delta_expected_Colorimetry_8 | recovery_pct_Colorimetry_8 | rel_error_Colorimetry_8
L | CIELAB | calibration_plus_stdadd | 3 | 0.41981 | slope_agreement^2 * sqrt(R2_cal * R2_std) * (1/LOQ) | calibration_plus_stdadd | 0.97997 | 0.92658 | -3.9186 | -2.2402 | 0.57169 | 0.57169 | 0.42831 | -528.9144 | 35.1541 | -528.9144 | 35.1541 | standard_addition | 0.22256 | 0.74187 | 13.4795 | 2 | 0 | -528.9144 | Colorimetry 4% | 1 | 97 | 2 | -641.1872 | -738.1872 | -661.0177 | -7.6102 | Colorimetry 8% | 2 | 98 | 2 | -416.6417 | -514.6417 | -425.1446 | -5.2514
DeltaL | DeltaCIELAB | calibration_plus_stdadd | 3 | 0.41981 | slope_agreement^2 * sqrt(R2_cal * R2_std) * (1/LOQ) | calibration_plus_stdadd | 0.97997 | 0.92658 | -3.9186 | -2.2402 | 0.57169 | 0.57169 | 0.42831 | 121.7918 | 13.2943 | 121.7918 | 13.2943 | standard_addition | 0.22256 | 0.74187 | 13.4795 | 2 | 0 | 121.7918 | Colorimetry 4% | 1 | 97 | 2 | 107.8186 | 10.8186 | 111.1532 | 0.11153 | Colorimetry 8% | 2 | 98 | 2 | 135.765 | 37.765 | 138.5358 | 0.38536
PAbs_Green | RGB | calibration_plus_stdadd | 3 | 0.37331 | slope_agreement^2 * sqrt(R2_cal * R2_std) * (1/LOQ) | calibration_only | 0.99525 | 0.89547 | 0.068738 | 0.036505 | 0.53108 | 0.53108 | 0.46892 | 143.8905 | 15.4878 | 143.8905 | 15.4878 | standard_addition | 0.21397 | 0.71325 | 14.0203 | 2 | 0 | 143.8905 | Colorimetry 4% | 1 | 97 | 2 | 129.181 | 32.181 | 133.1763 | 0.33176 | Colorimetry 8% | 2 | 98 | 2 | 158.6 | 60.6 | 161.8367 | 0.61837
DeltaE | DeltaCIELAB | calibration_plus_stdadd | 3 | 0.1667 | slope_agreement^2 * sqrt(R2_cal * R2_std) * (1/LOQ) | calibration_plus_stdadd | 0.98439 | 0.79924 | 7.3373 | 3.6254 | 0.4941 | 0.4941 | 0.5059 | 135.1129 | 26.2841 | 135.1129 | 26.2841 | standard_addition | 0.38972 | 1.2991 | 7.6979 | 2 | 0 | 135.1129 | Colorimetry 4% | 1 | 97 | 2 | 126.6483 | 29.6483 | 130.5653 | 0.30565 | Colorimetry 8% | 2 | 98 | 2 | 143.5775 | 45.5775 | 146.5077 | 0.46508
PAbs_Blue | RGB | calibration_plus_stdadd | 3 | 0.16378 | slope_agreement^2 * sqrt(R2_cal * R2_std) * (1/LOQ) | calibration_only | 0.86643 | 0.68558 | 0.0073746 | 0.0071625 | 0.9348 | 0.97123 | 0.065867 | 198.7774 | 56.6265 | 198.7774 | 56.6265 | standard_addition | 1.2337 | 4.1122 | 2.4318 | 2 | 0 | 198.7774 | Colorimetry 4% | 1 | 97 | 2 | 217.0987 | 120.0987 | 223.8131 | 1.2381 | Colorimetry 8% | 2 | 98 | 2 | 180.4562 | 82.4562 | 184.139 | 0.84139
Deltaa | DeltaCIELAB | calibration_plus_stdadd | 3 | 0.1211 | slope_agreement^2 * sqrt(R2_cal * R2_std) * (1/LOQ) | calibration_plus_stdadd | 0.97759 | 0.73036 | 4.2507 | 2.1251 | 0.49994 | 0.49994 | 0.50006 | 133.0148 | 28.9065 | 133.0148 | 28.9065 | standard_addition | 0.52318 | 1.7439 | 5.7341 | 2 | 0 | 133.0148 | Colorimetry 4% | 1 | 97 | 2 | 141.4997 | 44.4997 | 145.8759 | 0.45876 | Colorimetry 8% | 2 | 98 | 2 | 124.5299 | 26.5299 | 127.0713 | 0.27071
a | CIELAB | calibration_plus_stdadd | 3 | 0.1211 | slope_agreement^2 * sqrt(R2_cal * R2_std) * (1/LOQ) | calibration_plus_stdadd | 0.97759 | 0.73036 | 4.2507 | 2.1251 | 0.49994 | 0.49994 | 0.50006 | 107.2842 | 24.8485 | 107.2842 | 24.8485 | standard_addition | 0.52318 | 1.7439 | 5.7341 | 2 | 0 | 107.2842 | Colorimetry 4% | 1 | 97 | 2 | 112.0304 | 15.0304 | 115.4953 | 0.15495 | Colorimetry 8% | 2 | 98 | 2 | 102.5381 | 4.5381 | 104.6307 | 0.046307
DeltaE_chroma | DeltaCIELAB | calibration_plus_stdadd | 3 | 0.10086 | slope_agreement^2 * sqrt(R2_cal * R2_std) * (1/LOQ) | calibration_plus_stdadd | 0.97495 | 0.68501 | 6.284 | 2.8477 | 0.45317 | 0.45317 | 0.54683 | 143.1743 | 34.0365 | 143.1743 | 34.0365 | standard_addition | 0.49921 | 1.664 | 6.0095 | 2 | 0 | 143.1743 | Colorimetry 4% | 1 | 97 | 2 | 136.9956 | 39.9956 | 141.2325 | 0.41233 | Colorimetry 8% | 2 | 98 | 2 | 149.353 | 51.353 | 152.401 | 0.52401
PAbs_Red | RGB | calibration_plus_stdadd | 3 | 0.098527 | slope_agreement^2 * sqrt(R2_cal * R2_std) * (1/LOQ) | calibration_only | 0.98393 | 0.78301 | 0.029967 | 0.012891 | 0.43018 | 0.43018 | 0.56982 | 317.3561 | 60.8032 | 317.3561 | 60.8032 | standard_addition | 0.49458 | 1.6486 | 6.0657 | 2 | 0 | 317.3561 | Colorimetry 4% | 1 | 97 | 2 | 200.5534 | 103.5534 | 206.7561 | 1.0676 | Colorimetry 8% | 2 | 98 | 2 | 434.1587 | 336.1587 | 443.0191 | 3.4302
b | CIELAB | calibration_plus_stdadd | 3 | 0.078627 | slope_agreement^2 * sqrt(R2_cal * R2_std) * (1/LOQ) | calibration_plus_stdadd | 0.96941 | 0.6306 | -4.6445 | -1.8607 | 0.40064 | 0.40064 | 0.59936 | 180.7671 | 50.1849 | 180.7671 | 50.1849 | standard_addition | 0.47883 | 1.5961 | 6.2653 | 2 | 0 | 180.7671 | Colorimetry 4% | 1 | 97 | 2 | 157.0616 | 60.0616 | 161.9192 | 0.61919 | Colorimetry 8% | 2 | 98 | 2 | 204.4727 | 106.4727 | 208.6456 | 1.0865
Deltab | DeltaCIELAB | calibration_plus_stdadd | 3 | 0.078627 | slope_agreement^2 * sqrt(R2_cal * R2_std) * (1/LOQ) | calibration_plus_stdadd | 0.96941 | 0.6306 | -4.6445 | -1.8607 | 0.40064 | 0.40064 | 0.59936 | 160.5214 | 45.7892 | 160.5214 | 45.7892 | standard_addition | 0.47883 | 1.5961 | 6.2653 | 2 | 0 | 160.5214 | Colorimetry 4% | 1 | 97 | 2 | 137.0836 | 40.0836 | 141.3233 | 0.41323 | Colorimetry 8% | 2 | 98 | 2 | 183.9593 | 85.9593 | 187.7136 | 0.87714

## 11_CIELAB_FITTING
Channel | FitType | ID | DF | n_points | m | q | R2 | RMSE | LOD | LOQ | C0 | C0_sd | sigma_cal | sigma_source | SNR | beta_k | bias_index_k
DeltaL | Calibration |  |  | 11 | -3.9186 | 1.2237 | 0.97997 | 0.8569 | 0.22256 | 0.74187 |  |  | 0.29071 | median_calibration_sd | 13.4795
DeltaL | StdAdd | 1 | 25 | 11 | -2.6702 | -11.5159 | 0.96082 | 0.87246 |  |  | 73.4697 | 2.5454 |  | unavailable
DeltaL | StdAdd | 2 | 12.5 | 11 | -1.8103 | -19.6616 | 0.89234 | 0.99075 |  |  | 62.719 | 1.7374 |  | unavailable
Deltaa | Calibration |  |  | 11 | 4.2507 | 2.072 | 0.97759 | 1.0227 | 0.52318 | 1.7439 |  |  | 0.7413 | median_calibration_sd | 5.7341
Deltaa | StdAdd | 1 | 25 | 11 | 2.545 | 14.4048 | 0.82463 | 1.7438 |  |  | 84.7197 | 5.6862 |  | unavailable
Deltaa | StdAdd | 2 | 12.5 | 11 | 1.7052 | 16.9877 | 0.63608 | 1.7703 |  |  | 49.9553 | 2.7718 |  | unavailable
Deltab | Calibration |  |  | 11 | -4.6445 | -2.6171 | 0.96941 | 1.3174 | 0.47883 | 1.5961 |  |  | 0.7413 | median_calibration_sd | 6.2653
Deltab | StdAdd | 1 | 25 | 11 | -2.5028 | -13.7234 | 0.78953 | 1.9323 |  |  | 73.87 | 6.0201 |  | unavailable
Deltab | StdAdd | 2 | 12.5 | 11 | -1.2187 | -17.9356 | 0.47168 | 1.7345 |  |  | 48.2714 | 2.6102 |  | unavailable
DeltaE_ab | Calibration |  |  | 11 | 7.3373 | 2.3778 | 0.98439 | 1.4524 | 0.38972 | 1.2991 |  |  | 0.95316 | median_calibration_sd | 7.6979
DeltaE_ab | StdAdd | 1 | 25 | 11 | 4.4963 | 22.778 | 0.87593 | 2.5784 |  |  | 77.6102 | 5.0336 |  | unavailable
DeltaE_ab | StdAdd | 2 | 12.5 | 11 | 2.7544 | 31.6381 | 0.72255 | 2.4868 |  |  | 53.8995 | 2.3472 |  | unavailable
DeltaE_ab_chroma | Calibration |  |  | 11 | 6.284 | 3.2887 | 0.97495 | 1.6075 | 0.49921 | 1.664 |  |  | 1.0457 | median_calibration_sd | 6.0095
DeltaE_ab_chroma | StdAdd | 1 | 25 | 11 | 3.6233 | 19.855 | 0.81154 | 2.5794 |  |  | 78.9903 | 5.657 |  | unavailable
DeltaE_ab_chroma | StdAdd | 2 | 12.5 | 11 | 2.0721 | 24.7584 | 0.55848 | 2.4574 |  |  | 49.2489 | 2.6039 |  | unavailable
DeltaL | Calibration |  |  | 11 | -3.9186 | 1.2237 | 0.97997 | 0.8569 | 0.22256 | 0.74187 |  |  | 0.29071 | median_calibration_sd | 13.4795
DeltaL | StdAdd | 1 | 25 | 11 | -2.6702 | -11.5159 | 0.96082 | 0.87246 |  |  | 73.4697 | 2.5454 |  | unavailable
DeltaL | StdAdd | 2 | 12.5 | 11 | -1.8103 | -19.6616 | 0.89234 | 0.99075 |  |  | 62.719 | 1.7374 |  | unavailable

## 12_LEGENDS
Term | Meaning | Formula | Unit | Where used | Notes
area | Accepted background-mask area | number of accepted pixels in the background sample mask | pixels | 02_BG_SAMPLES | Area after model-based and robust statistical background-mask filtering.
Associated_Wells | Four wells surrounding an inter-well background cell | well(r,c)-well(r,c+1)-well(r+1,c)-well(r+1,c+1) | well labels | 02_BG_SAMPLES | Clarifies that BG samples are inter-well regions rather than wells.
B_*/G_*/R_* | Robust statistics of raw well-channel intensities in OpenCV order | computed over retained well ROI pixels; suffix = mean, median, sd, p10, p25, p50, p75, p90 or iqr | raw image intensity | 04_WELL_ROBUST_STATS | OpenCV internal channel order is B,G,R.
B_bg/G_bg/R_bg | Predicted local raw background at a well | 2D background surface evaluated at well center | raw image intensity | 03_BG_WELL_FIT | OpenCV internal channel order is B,G,R.
B_med/G_med/R_med | Median raw OpenCV-channel value in accepted BG-mask pixels | median over accepted background-mask pixels | raw image intensity | 02_BG_SAMPLES | OpenCV internal channel order is B,G,R.
beta_k/bias_index_k | Per-fit slope ratio and relative slope-bias index | beta_k = m_std/m_cal; bias_index_k = ¦beta_k - 1¦ | dimensionless | 11_CIELAB_FITTING
beta_mean | Mean standard-addition/calibration slope ratio | mean(m_std / m_cal) | dimensionless | 10_METHOD_COMPARISON | 1 indicates equal slopes on average.
beta_mean/bias_index_mean | Mean slope ratio and mean relative slope bias | beta_mean = mean(m_std/m_cal); bias_index_mean = mean(¦m_std/m_cal - 1¦) | dimensionless | 10_METHOD_COMPARISON
BG_Cell_Row/BG_Cell_Col | Inter-well background-cell row and column index | 0-based index of the inter-well cell, not a well coordinate | index | 02_BG_SAMPLES | BG_Cell_Row=0 and BG_Cell_Col=0 are surrounded by A1, A2, B1 and B2.
BG_Red_raw/BG_Green_raw/BG_Blue_raw | Predicted local raw background at a well | 2D background surface evaluated at well center | raw image intensity | 03_BG_WELL_FIT | Exported in standard RGB order; used to compute per-well MeanBG after linearization.
BG_STAT_MASK | Binary diagnostic mask of accepted inter-well background pixels | white pixels = accepted for background sampling; black pixels = rejected | image | BG_STAT_MASK.png, RAW_DATA_DETAILS_CAPTION.txt | Useful for auditability and troubleshooting of background sampling.
bias_index_mean | Mean relative slope bias | mean(¦m_std/m_cal - 1¦) | dimensionless | 10_METHOD_COMPARISON | 0 indicates no slope bias.
BrightExcessMeanGray | Mean excess brightness above the local bright threshold | mean(max(Gray - bright_threshold, 0)) over bright-excluded pixels | raw grayscale intensity | 04_WELL_ROBUST_STATS | Used with BrightExcludedFraction to form HighlightIndex.
BrightExcludedFraction | Fraction of core ROI rejected as bright | n_bright_excluded / n_core | dimensionless | 04_WELL_ROBUST_STATS | Part of optical quality diagnostics.
BrightExcludedFraction/BrightExcludedMeanGray/BrightExcessMeanGray/HighlightIndex | Highlight and bright-pixel QC descriptors | HighlightIndex = BrightExcludedFraction x BrightExcessMeanGray | mixed | 04_WELL_ROBUST_STATS, 08_EMPTY_WELLS
BrightExcludedMeanGray | Mean grayscale intensity of bright-excluded pixels | mean(Gray) over pixels excluded as overly bright | raw grayscale intensity | 04_WELL_ROBUST_STATS | Blank if no bright pixels were excluded.
C0/C0_sd | Estimated original concentration and associated SD | standard-addition or calibration projection depending on FitType | mM | 11_CIELAB_FITTING | Diagnostic for CIELAB/DeltaE fits.
C0_mean/C0_median/C0_sd_median | Summary of available original-concentration estimates | mean/median of C0 estimates and median of associated C0_sd values | mM | 10_METHOD_COMPARISON
C0_median/C0_sd_median | Median estimated original concentration and median associated SD | median over available standard-addition or unknown estimates | mM | 10_METHOD_COMPARISON | Reported for diagnostic comparison; expected/reference values are external checks.
Channel | Fitted diagnostic descriptor | L, a, b, DeltaL, Deltaa, Deltab, DeltaE or DeltaE_chroma | text | 11_CIELAB_FITTING | CIELAB/DeltaE descriptors are diagnostic, not primary quantitative outputs.
Channel/FitType/ID/DF/n_points/m/q/R2/RMSE | Diagnostic fitting identifiers and linear-fit descriptors | y = m x + q; R2 = 1 - SSE/SST; RMSE = sqrt(mean(residual^2)) | mixed | 11_CIELAB_FITTING
Code implementation | Main computational libraries and code provenance | custom Python code using NumPy, OpenCV, openpyxl and matplotlib | software provenance | DIAGNOSTICS.xlsx and diagnostic PNG files | Sources/libraries: OpenCV for masks and BGR-to-Lab conversion; NumPy for numerical fitting; openpyxl for xlsx export; matplotlib for PNG figures.
CommonFactorsN | Number of common factors used in the score | 4 for R2_cal, R2_std, SlopeAgreement and LOQ; 1 for calibration-only or stdadd-only fallbacks | integer | 10_METHOD_COMPARISON | Used to avoid comparing scores obtained from different formulas.
ComparableGroup | Set of methods scored with the same formula | calibration_plus_stdadd, calibration_only, stdadd_only, or not_ranked | text | 10_METHOD_COMPARISON | Scores are directly comparable only within the same group.
corr_col/corr_row | Correlation with column/row position | Pearson correlation | dimensionless | 09_SPATIAL_DIAGNOSTICS | Diagnostic only.
D_warning/D_critical | Floor-geometry warning and critical flags | rule-based thresholds on floor/mouth geometry descriptors | 0/1 | 05_GEOMETRY_QC | Critical is stricter than warning.
dataset | Dataset used for spatial trend | unknown or empty | text | 09_SPATIAL_DIAGNOSTICS | Spatial diagnostics are descriptive and do not alter results.
Dataset/Status/Applicability/Reason | Applicability status for optional spatial diagnostics | applied or not_applied with reason text | text | 09_SPATIAL_DIAGNOSTICS | The sheet is retained even when spatial diagnostics cannot be applied.
DeltaE_ab | Total CIELAB color difference | sqrt(DeltaL^2 + Deltaa^2 + Deltab^2) | dimensionless | REPORT, 10_METHOD_COMPARISON, 11_CIELAB_FITTING, FIGURE_CIELAB_DELTAE.png | Reference source is reported in CIELAB_ref_source in the main report.
DeltaE_ab_chroma | Chromatic CIELAB difference | sqrt(Deltaa^2 + Deltab^2) | dimensionless | REPORT, 10_METHOD_COMPARISON, 11_CIELAB_FITTING, FIGURE_CIELAB_DELTAE.png | Excludes the lightness term DeltaL.
DeltaL/Deltaa/Deltab/DeltaE_ab/DeltaE_ab_chroma | CIELAB difference descriptors | DeltaL = L - L_ref; Deltaa = a - a_ref; Deltab = b - b_ref; DeltaE_ab = sqrt(DeltaL^2 + Deltaa^2 + Deltab^2); DeltaE_ab_chroma = sqrt(Deltaa^2 + Deltab^2) | dimensionless | FIGURE_CIELAB_DELTAE.png, RAW_DATA_DETAILS_CAPTION.txt, 10_METHOD_COMPARISON, 11_CIELAB_FITTING | Reference: CIE 1976 L*a*b* color-difference form.
estimate_for_expected_*/delta_expected_*/recovery_pct_*/rel_error_* | External-reference comparison metrics | delta = estimate - reference; recovery = 100 x estimate/reference; relative error = 100 x (estimate - reference)/reference | mM or % | METHOD_COMPARISON.png, 10_METHOD_COMPARISON
Estimate_source | Source of the representative concentration estimate | standard_addition, unknown_from_calibration, epsilon or unavailable | text | 10_METHOD_COMPARISON | Identifies how Estimate_value was obtained.
Estimate_value/Estimate_sd | Representative diagnostic concentration estimate and associated SD | derived according to Estimate_source | mM | 10_METHOD_COMPARISON
Estimate_value/Estimate_sd/Estimate_source | Representative diagnostic concentration estimate, SD and source | derived according to Estimate_source | mM and text | METHOD_COMPARISON.png, 10_METHOD_COMPARISON
expected_label_*/expected_id_*/expected_value_*/expected_sd_* | External reference metadata | user/configurator input | text or mM | METHOD_COMPARISON.png, 10_METHOD_COMPARISON | External reference values are checks only and are not part of Score.
Family | Method family | RGB, CIELAB, DeltaCIELAB or other | text | 10_METHOD_COMPARISON | Used to separate method families.
FitType | Type of fit row | Calibration, StdAdd, UnknownFromCal, UnknownOnly | text | 11_CIELAB_FITTING | Same convention as main fitting output.
floor_source/local_pitch_px/px_per_mm/cyl_r_bg | Geometry-source and local scale descriptors | reported source, local pitch, pixel/mm scale and background distance-transform radius | mixed | 05_GEOMETRY_QC, 06_WELL_BOTTOM
floor_to_mouth_r_ratio | Relative floor radius | floor_r / mouth_r | dimensionless | 05_GEOMETRY_QC | Used to flag abnormal floor geometry.
Geometry and epsilon/path-length quantification | Assumption used when epsilon-based unknown quantification is enabled | l_cm = (volume_uL / well_bottom_area_mm2) / 10; C_M = PAbs / (epsilon x l_cm) | cm and M | RESULTS_CAPTION.txt, RAW_DATA_DETAILS_CAPTION.txt, REPORT | Assumes ANSI/SLAS-compatible flat-bottom microplate geometry; non-flat or non-certified geometries require separate validation.
Gray_* | Robust statistics of grayscale intensity | computed over retained ROI pixels; suffix = mean, median, sd, p10, p25, p50, p75, p90 or iqr | raw image intensity | 04_WELL_ROBUST_STATS | Diagnostic descriptor used for optical QC.
highlight_fraction_roi/highlight_fraction_core | Fraction of very bright pixels in ROI/core | fraction of pixels with grayscale above the highlight threshold | dimensionless | 04_WELL_ROBUST_STATS | Optical QC descriptor for highlights/specular artifacts.
HighlightIndex | Combined highlight severity index | BrightExcludedFraction x BrightExcessMeanGray | gray-level weighted fraction | 04_WELL_ROBUST_STATS | Higher values indicate more severe bright artifacts.
ImageWarning/WarningReason | Well-level optical QC warning and triggering reason | rule-based flags from highlights, trimming and intensity SD | 0/1 and text | 04_WELL_ROBUST_STATS | WarningReason lists the triggering conditions.
IRLS | Iteratively reweighted least-squares robust linear regression with residual-based weights | minimize sum_i w_i (y_i - (m x_i + q))^2; w_i is updated iteratively from residual magnitude | dimensionless | METHOD_COMPARISON.png, FIGURE_CIELAB_DELTAE.png, 10_METHOD_COMPARISON, 11_CIELAB_FITTING | Reference: Huber, P. J. (1964), Robust Estimation of a Location Parameter. Implementation: custom NumPy-based IRLS using repeated least-squares solves.
key/value | Plate-geometry metadata key and value | embedded plate-geometry database entry | mixed | 07_PLATE_GEOMETRY | Geometry is assumed to be ANSI/SLAS-compatible flat-bottom geometry unless stated otherwise.
L_*/a_*/b_* | Robust statistics of CIELAB coordinates | computed from OpenCV CIELAB conversion over retained ROI pixels; suffix = mean, median, sd, p10, p25, p50, p75, p90 or iqr | dimensionless | 04_WELL_ROBUST_STATS | CIELAB values are diagnostic descriptors derived from RGB image data.
local_pitch_px / px_per_mm / cyl_r_bg | Local geometry scale descriptors | local plate pitch, pixel-to-mm conversion and cylindrical background radius | pixels or pixels/mm | 06_WELL_BOTTOM | Technical geometry diagnostics.
LOD/LOQ | Detection and quantification limits | LOD = 3 sigma_cal / ¦m¦; LOQ = 10 sigma_cal / ¦m¦ | mM | 11_CIELAB_FITTING, 10_METHOD_COMPARISON | Diagnostic for CIELAB/DeltaE fits.
m/q/R2/RMSE | Linear-fit parameters and fit quality | y = m x + q; R2 coefficient of determination; RMSE root-mean-square error | descriptor units | 11_CIELAB_FITTING | Diagnostic CIELAB/DeltaE fitting only.
m_cal/m_std_mean | Calibration slope and mean standard-addition slope | slope from linear fit y = m x + q | signal/mM | 10_METHOD_COMPARISON, 11_CIELAB_FITTING | Used to compute slope agreement.
MeanBG_* | Linearized local background intensity | background surface value after gamma linearization | dimensionless | REPORT | The asterisk denotes Red, Green or Blue.
MeanBG_Red/MeanBG_Green/MeanBG_Blue | Linearized local background intensity for each RGB channel | local background value after gamma linearization | dimensionless | 08_EMPTY_WELLS
MeanW_* | Linearized median well intensity | median well ROI intensity after gamma linearization | dimensionless | REPORT | The asterisk denotes Red, Green or Blue.
MeanW_Red/MeanW_Green/MeanW_Blue | Linearized median well intensity for each RGB channel | median ROI channel intensity after gamma linearization | dimensionless | 08_EMPTY_WELLS
Method | Compared analytical descriptor | PAbs_*, L/a/b, Delta* or DeltaE* | text | 10_METHOD_COMPARISON | PAbs_* are RGB; CIELAB and DeltaCIELAB are derived diagnostics.
Method/Family/ComparableGroup/CommonFactorsN/RankMode | Method-comparison identifiers and comparable-score grouping | text labels and number of factors defining score comparability | mixed | 10_METHOD_COMPARISON, METHOD_COMPARISON.png | Scores are directly comparable only within the same ComparableGroup.
mouth_* / floor_* | Mouth/floor circle center, radius and score descriptors | projected or refined circle parameters | pixels or score | 06_WELL_BOTTOM | Used for ROI and geometry QC.
n/intercept/slope_col/slope_row/R2/corr_col/corr_row | Spatial-trend fit descriptors | linear trend of the diagnostic signal versus row/column position | mixed | 09_SPATIAL_DIAGNOSTICS | Diagnostic only; does not alter quantitative results.
n_roi/n_core/n_used | Pixel counts used during well ROI filtering | ROI pixels, core pixels and retained pixels | pixels | 04_WELL_ROBUST_STATS | Used to audit well-level pixel filtering.
PAbs_* | RGB pseudo-absorbance | log10(MeanBG_*/MeanW_*) | dimensionless | REPORT, 10_METHOD_COMPARISON | Primary RGB analytical descriptor.
PAbs_Red/PAbs_Green/PAbs_Blue | RGB pseudo-absorbance channels | log10(MeanBG_channel / MeanW_channel) | dimensionless | 08_EMPTY_WELLS, 10_METHOD_COMPARISON, METHOD_COMPARISON.png | Primary quantitative RGB descriptor; diagnostic sheets may include it for comparison.
Purple_* | Robust statistics of the internal purple-color index | Purple = 0.5 x (Red + Blue) - Green; suffix = mean, median, sd, p10, p25, p50, p75, p90 or iqr | raw image intensity | 04_WELL_ROBUST_STATS | Diagnostic color index, not the primary quantitative signal.
Purpose | Short description of the sheet role | free-text description | text | 01_CONTENTS
R2_cal/R2_std_mean | Calibration and mean standard-addition coefficient of determination | R2 from linear fits; R2_std_mean is averaged over standard-addition curves | dimensionless | 10_METHOD_COMPARISON, 11_CIELAB_FITTING | Higher values indicate better linear fit quality.
R2_cal/R2_std_mean/m_cal/m_std_mean/SlopeAgreement | Fit-quality and slope-agreement descriptors for method comparison | SlopeAgreement = min(¦m_cal¦, ¦m_std_mean¦) / max(¦m_cal¦, ¦m_std_mean¦) | mixed | 10_METHOD_COMPARISON, METHOD_COMPARISON.png
RankMode | Data basis available for ranking | calibration_plus_stdadd, calibration_only, stdadd_only, or unavailable | text | 10_METHOD_COMPARISON | Only rows with the same comparable group should be directly compared.
Red_*/Green_*/Blue_* | Robust statistics of raw well-channel intensities | computed over retained well ROI pixels; *=mean, median, sd, p10, p25, p50, p75, p90, iqr | raw image intensity | 04_WELL_ROBUST_STATS | Exported in standard RGB order.
Red_median_raw/Green_median_raw/Blue_median_raw | Median raw channel value in accepted BG-mask pixels | median over accepted mask pixels | raw image intensity | 02_BG_SAMPLES | Exported in standard RGB order; internally OpenCV stores images in BGR order.
reference_*/delta_reference_*/recovery_pct_* | External reference checks used in diagnostic comparison | delta = estimate - reference; recovery = 100 x estimate/reference | mM or % | 10_METHOD_COMPARISON, METHOD_COMPARISON.png | Reference values are external checks and are not part of Score.
RGB to CIELAB | Diagnostic conversion from RGB/BGR to CIE L*a*b* | OpenCV BGR->Lab for 8-bit sRGB-like images; reference form: linearized sRGB -> XYZ using the standard sRGB-to-XYZ matrix and D65 reference white -> CIE L*a*b* | dimensionless | FIGURE_CIELAB_DELTAE.png, RAW_DATA_DETAILS_CAPTION.txt, 04_WELL_ROBUST_STATS, 11_CIELAB_FITTING | References: OpenCV color conversions; IEC 61966-2-1:1999 sRGB; CIE 1976 L*a*b*; CIE standard illuminant D65.
Row/Col/Well | Human-readable well position | Row is A-based; Col is 1-based; Well = Row + Col | well label | 03_BG_WELL_FIT, 04_WELL_ROBUST_STATS, 05_GEOMETRY_QC, 06_WELL_BOTTOM, 08_EMPTY_WELLS | Used for true well-level records, not for inter-well BG cells.
Score | Common ranking score | for calibration+standard addition: SlopeAgreement^2 x sqrt(R2_cal x R2_std) x (1/LOQ); fallback groups use the formula stated in ScoreFormula | 1/mM | 10_METHOD_COMPARISON, METHOD_COMPARISON.png | Expected/reference values, SNR and clipping are excluded.
Score/ScoreFormula | Common method-ranking score and formula descriptor | SlopeAgreement^2 x sqrt(R2_cal x R2_std_mean) x (1/LOQ) for calibration_plus_stdadd rows | 1/mM and text | 10_METHOD_COMPARISON, METHOD_COMPARISON.png | Expected/reference values, SNR and clipping are not part of Score.
ScoreFormula | Formula used to compute Score | text descriptor | text | 10_METHOD_COMPARISON | Documents which score formula was applied to the row.
Sheet | Workbook sheet name | worksheet name | text | 01_CONTENTS
shift_frac_of_mouth_r | Relative mouth-to-floor center shift | shift_px / mouth_r | dimensionless | 05_GEOMETRY_QC | Large values indicate poor floor/mouth alignment.
shift_px | Mouth-to-floor center shift | Euclidean distance between fitted mouth and floor centers | pixels | 06_WELL_BOTTOM, 05_GEOMETRY_QC | Large values indicate poor floor/mouth alignment.
shift_px/shift_frac_of_mouth_r/floor_to_mouth_r_ratio/floor_to_mouth_area_ratio | Mouth-to-floor alignment descriptors | shift distance and relative floor/mouth radius or area ratios | pixels or dimensionless | 05_GEOMETRY_QC, 06_WELL_BOTTOM
sigma_cal/sigma_source/SNR | Calibration noise estimate, its source and slope-to-noise ratio | SNR = ¦m¦/sigma_cal | mixed | 11_CIELAB_FITTING, 10_METHOD_COMPARISON | SNR is diagnostic and is not part of the common Score.
slope_col/slope_row | Spatial trend coefficients | linear trend versus column and row | signal/index | 09_SPATIAL_DIAGNOSTICS | Used to identify plate-position effects.
SlopeAgreement | Agreement between calibration and standard-addition slopes | min(¦m_cal¦, ¦m_std¦) / max(¦m_cal¦, ¦m_std¦) | dimensionless | 10_METHOD_COMPARISON | 1 indicates identical slope magnitude.
Status/Applicability/Reason | Applicability statement for optional spatial diagnostics | applied/not_applied with reason text | text | 09_SPATIAL_DIAGNOSTICS | The sheet is retained even when spatial diagnostics cannot be applied.
used_fraction/UsedFraction | Fraction of ROI core pixels retained after filtering | n_used / n_core | dimensionless | 04_WELL_ROBUST_STATS, 08_EMPTY_WELLS
UsedFraction | Fraction of core ROI retained | n_used / n_core | dimensionless | 04_WELL_ROBUST_STATS | Low values indicate strong trimming.
x/y | Image coordinates of a background sample or well center | pixel coordinate in the analyzed image | pixels | 02_BG_SAMPLES, 03_BG_WELL_FIT | Coordinate system follows the image array.