## 01_CONTENTS
Sheet | Purpose
01_CONTENTS | Index of primary results workbook sheets.
02_METADATA | Image-level metadata and rule-based image QC used to audit the analysis.
03_OVERVIEW | Final quantitative summary, selected method, reliability and external reference checks when provided.
04_RAW | Well-level analytical values used by the fitting pipeline.
05_REPLICATES_MEAN | Replicate-group medians, robust SDs and group QC flags.
06_FITTING | Calibration, standard-addition and unknown/CRM fit results.
07_METHOD_COMPARISON | Method ranking using only common score factors; expected values are external checks only.
08_LEGENDS | Definitions for all fields reported in this workbook and primary RGB figure.

## 02_METADATA
Field | Value | Notes
Analysis image source | original | original or in-memory resized image used for all measurements
Original image path | C:/Users/prova/OneDrive - unifi.it/Desktop/cromo/Piastra viola/U1.jpg | source file selected by the user
Original image size | 1787 x 1203 px | width x height in pixels, when known
Analysis image size | 1787 x 1203 px | width x height in pixels actually analyzed
Resize scale | 1 | analysis_width / original_width, when a resized image is used
Initial image QC | FAIL | Rule-based result: FAIL if destructive=True; WARNING if quality_warning=True or borderline=True; otherwise OK. No image correction is applied.
Initial image QC class | non_correctable | non_correctable if destructive=True; quality_warning if flatfield_span > 0.18 and no destructive defect; usable_with_warnings if only borderline defects; good if no flags.
Initial image QC messages | saturation; specular reflections | Messages list the threshold rules triggered by the measured fields below.
Image QC decision rules | thresholds | destructive = dead_channel OR saturation_fraction>0.003 OR saturation_all_channels_fraction>0.0008 OR specular_fraction>0.003 OR max_side<900 OR approx_roi_pixels<120; borderline = saturation_fraction>0.0005 OR specular_fraction>0.0005 OR blur_score<35; quality_warning = flatfield_span>0.18.
Flat-field span | 0.1332 | Flatfield_span = (P95 - P5) / median of the slow illumination field estimated from the grayscale image. Larger values indicate stronger illumination/background gradient.
Specular fraction | 0.014841 | fraction of bright low-saturation pixels
Saturation fraction | 0.017505 | fraction of pixels with at least one saturated channel
Dead channel | 0 | 1 if a channel is near-dead by mean/variance checks
Approx well pitch | 124.4775 | median distance between adjacent well centers, in pixels
Approx ROI pixels/well | 4441 | median number of pixels used in the final floor ROI mask per well
Blur score | 76.678 | variance of Laplacian of the analysis image; lower values indicate stronger blur

## 03_OVERVIEW
Field | Value
Quantification | available
Reliability score | 17
Confidence class | LOW
Reliability note | concentration estimated from PAbs with epsilon because no calibration was available; no quantitative method was ranked as available; epsilon-based quantification without calibration has intrinsically limited reliability
Empty-well QC status | ok
Configured epsilon (M-1 cm-1) | 180
Liquid volume per well (uL) | 200
Calculated path length (cm) | 0.60644
Calculated liquid height (mm) | 6.0644
Path-length source | calculated_from_volume_and_flat_bottom_geometry
Plate geometry for path length | 96-well standard F-bottom / flat-bottom
Plate geometry assumption | ANSI/SLAS-compatible flat-bottom geometry; vendor dimensions may vary by plate family
Selected method from rank | L
Selected family | CIELAB
Ranking mode
Selected method score | 0
Status notes | epsilon quantification used; empty-well QC: ok; epsilon configured; path length calculated_from_volume_and_flat_bottom_geometry

## 04_RAW
Row | Col | Well | ID | Type | Conc | DF | MeanW_Red | MeanW_Green | MeanW_Blue | MeanBG_Red | MeanBG_Green | MeanBG_Blue | SignalT_Red | SignalT_Green | SignalT_Blue | PAbs_Red | PAbs_Green | PAbs_Blue | L | a | b | DeltaL | Deltaa | Deltab | DeltaE_ab | DeltaE_ab_chroma | CIELAB_ref_source | ImageWarning
A | 1 | A1 | 1 | U |  | 1 | 0.45908 | 0.35374 | 0.55423 | 0.61443 | 0.52456 | 0.74673 | 0.74716 | 0.67436 | 0.7422 | 0.12658 | 0.17111 | 0.12948 | 68.2353 | 14 | -16 | 0.39216 | -3 | 3 | 4.2607 | 4.2426 | global_median | 0
A | 2 | A2 | 1 | U |  | 1 | 0.44787 | 0.33446 | 0.5605 | 0.60755 | 0.53192 | 0.72261 | 0.73717 | 0.62877 | 0.77566 | 0.13243 | 0.20151 | 0.11033 | 67.0588 | 16 | -18 | -0.78432 | -1 | 1 | 1.6171 | 1.4142 | global_median | 0
A | 3 | A3 | 1 | U |  | 1 | 0.44787 | 0.33446 | 0.5605 | 0.60152 | 0.53899 | 0.69727 | 0.74456 | 0.62053 | 0.80385 | 0.1281 | 0.20724 | 0.094827 | 67.0588 | 16 | -18 | -0.78432 | -1 | 1 | 1.6171 | 1.4142 | global_median | 0
A | 4 | A4 | 1 | U |  | 1 | 0.45346 | 0.33922 | 0.56681 | 0.59633 | 0.54574 | 0.67082 | 0.76041 | 0.62158 | 0.84495 | 0.11895 | 0.2065 | 0.073167 | 67.451 | 16 | -18 | -0.39216 | -1 | 1 | 1.4676 | 1.4142 | global_median | 0
A | 5 | A5 | 1 | U |  | 1 | 0.46474 | 0.35865 | 0.56681 | 0.59196 | 0.55218 | 0.64335 | 0.78509 | 0.64952 | 0.88103 | 0.10508 | 0.18741 | 0.055009 | 68.6274 | 14 | -16 | 0.78431 | -3 | 3 | 4.3145 | 4.2426 | global_median | 0
B | 1 | B1 | 1 | U |  | 1 | 0.47618 | 0.35374 | 0.57955 | 0.64973 | 0.56632 | 0.78859 | 0.73289 | 0.62463 | 0.73491 | 0.13496 | 0.20438 | 0.13376 | 68.6274 | 16 | -17 | 0.78431 | -1 | 2 | 2.3696 | 2.2361 | global_median | 0
B | 2 | B2 | 1 | U |  | 1 | 0.47044 | 0.34403 | 0.59894 | 0.6408 | 0.57238 | 0.76495 | 0.73415 | 0.60104 | 0.78298 | 0.13422 | 0.2211 | 0.10625 | 68.2353 | 17 | -20 | 0.39216 | 0 | -1 | 1.0741 | 1 | global_median | 0
B | 3 | B3 | 1 | U |  | 1 | 0.46474 | 0.34403 | 0.60548 | 0.63276 | 0.5781 | 0.74 | 0.73447 | 0.59509 | 0.81822 | 0.13403 | 0.22541 | 0.087128 | 68.2353 | 17 | -21 | 0.39216 | 0 | -2 | 2.0381 | 2 | global_median | 0
B | 4 | B4 | 1 | U |  | 1 | 0.45908 | 0.33922 | 0.59244 | 0.6256 | 0.58347 | 0.71383 | 0.73383 | 0.58139 | 0.82994 | 0.13441 | 0.23553 | 0.080951 | 67.8431 | 17 | -20 | 0 | 0 | -1 | 1 | 1 | global_median | 0
B | 5 | B5 | 1 | U |  | 1 | 0.45346 | 0.34403 | 0.56681 | 0.61929 | 0.58848 | 0.68654 | 0.73221 | 0.5846 | 0.8256 | 0.13536 | 0.23314 | 0.083228 | 67.8431 | 15 | -17 | 0 | -2 | 2 | 2.8284 | 2.8284 | global_median | 0
C | 1 | C1 | 1 | U |  | 1 | 0.47618 | 0.35865 | 0.59894 | 0.67325 | 0.59512 | 0.81319 | 0.70729 | 0.60265 | 0.73653 | 0.1504 | 0.21993 | 0.13281 | 69.0196 | 17 | -18 | 1.1765 | 0 | 1 | 1.544 | 1 | global_median | 0
C | 2 | C2 | 1 | U |  | 1 | 0.48777 | 0.36859 | 0.63878 | 0.6622 | 0.59962 | 0.7903 | 0.73659 | 0.61471 | 0.80827 | 0.13278 | 0.21133 | 0.092442 | 69.8039 | 17 | -21 | 1.9608 | 0 | -2 | 2.8008 | 2 | global_median | 0
C | 3 | C3 | 1 | U |  | 1 | 0.5114 | 0.37868 | 0.69408 | 0.65208 | 0.60373 | 0.76603 | 0.78426 | 0.62722 | 0.90607 | 0.10554 | 0.20258 | 0.042837 | 70.9804 | 19 | -23 | 3.1373 | 2 | -4 | 5.4628 | 4.4721 | global_median | 0
C | 4 | C4 | 1 | U |  | 1 | 0.47044 | 0.34886 | 0.61207 | 0.64287 | 0.60746 | 0.74047 | 0.73178 | 0.5743 | 0.82659 | 0.13562 | 0.24086 | 0.082709 | 68.6274 | 17 | -20 | 0.78431 | 0 | -1 | 1.2709 | 1 | global_median | 0
C | 5 | C5 | 1 | U |  | 1 | 0.45908 | 0.34403 | 0.57316 | 0.63456 | 0.6108 | 0.71371 | 0.72346 | 0.56324 | 0.80307 | 0.14058 | 0.24931 | 0.095244 | 67.8431 | 16 | -18 | 0 | -1 | 1 | 1.4142 | 1.4142 | global_median | 0
D | 1 | D1 | 1 | U |  | 1 | 0.47618 | 0.3636 | 0.60548 | 0.68425 | 0.60981 | 0.81961 | 0.69591 | 0.59626 | 0.73874 | 0.15744 | 0.22457 | 0.13151 | 69.0196 | 17 | -19 | 1.1765 | 0 | 0 | 1.1765 | 0 | global_median | 0
D | 2 | D2 | 1 | U |  | 1 | 0.50543 | 0.38891 | 0.7083 | 0.67107 | 0.61253 | 0.79771 | 0.75318 | 0.63493 | 0.88792 | 0.1231 | 0.19728 | 0.051627 | 71.3726 | 18 | -23 | 3.5294 | 1 | -4 | 5.4274 | 4.1231 | global_median | 0
D | 3 | D3 | 1 | U |  | 1 | 0.68002 | 0.50543 | 0.95737 | 0.65886 | 0.61484 | 0.77437 | 1.0321 | 0.82206 | 1.2363 | -0.01373 | 0.085096 | -0.092132 | 80 | 22 | -28 | 12.1569 | 5 | -9 | 15.9308 | 10.2956 | global_median | 1
D | 4 | D4 | 1 | U |  | 1 | 0.49362 | 0.37362 | 0.67305 | 0.6476 | 0.61673 | 0.74969 | 0.76223 | 0.6058 | 0.89777 | 0.11792 | 0.21767 | 0.046834 | 70.5882 | 18 | -22 | 2.7451 | 1 | -3 | 4.1875 | 3.1623 | global_median | 0
D | 5 | D5 | 1 | U |  | 1 | 0.46474 | 0.34886 | 0.58597 | 0.63726 | 0.61822 | 0.72376 | 0.72928 | 0.56431 | 0.80963 | 0.13711 | 0.24848 | 0.091714 | 68.2353 | 16 | -19 | 0.39216 | -1 | 0 | 1.0741 | 1 | global_median | 0
E | 1 | E1 | 1 | U |  | 1 | 0.47044 | 0.35374 | 0.58597 | 0.68239 | 0.60979 | 0.80762 | 0.6894 | 0.5801 | 0.72555 | 0.16153 | 0.23649 | 0.13933 | 68.6274 | 16 | -18 | 0.78431 | -1 | 1 | 1.6171 | 1.4142 | global_median | 0
E | 2 | E2 | 1 | U |  | 1 | 0.47044 | 0.35865 | 0.63204 | 0.66713 | 0.61059 | 0.78688 | 0.70517 | 0.58739 | 0.80322 | 0.15171 | 0.23107 | 0.095163 | 69.0196 | 17 | -21 | 1.1765 | 0 | -2 | 2.3204 | 2 | global_median | 0
E | 3 | E3 | 1 | U |  | 1 | 0.5114 | 0.39408 | 0.72992 | 0.65288 | 0.61097 | 0.76469 | 0.78329 | 0.64502 | 0.95453 | 0.10608 | 0.19043 | 0.020211 | 71.7647 | 19 | -24 | 3.9216 | 2 | -5 | 6.6617 | 5.3852 | global_median | 0
E | 4 | E4 | 1 | U |  | 1 | 0.47618 | 0.35374 | 0.63204 | 0.63962 | 0.61091 | 0.74113 | 0.74447 | 0.57904 | 0.85281 | 0.12815 | 0.23729 | 0.069148 | 69.0196 | 18 | -21 | 1.1765 | 1 | -2 | 2.5267 | 2.2361 | global_median | 0
E | 5 | E5 | 1 | U |  | 1 | 0.45908 | 0.34403 | 0.58597 | 0.6273 | 0.61044 | 0.71629 | 0.73183 | 0.56357 | 0.81807 | 0.13559 | 0.24905 | 0.08721 | 68.2353 | 16 | -19 | 0.39216 | -1 | 0 | 1.0741 | 1 | global_median | 0
F | 1 | F1 | 1 | U |  | 1 | 0.44787 | 0.33446 | 0.5605 | 0.66774 | 0.59506 | 0.77766 | 0.67073 | 0.56205 | 0.72075 | 0.17345 | 0.25022 | 0.14222 | 67.0588 | 16 | -18 | -0.78432 | -1 | 1 | 1.6171 | 1.4142 | global_median | 0
F | 2 | F2 | 1 | U |  | 1 | 0.44232 | 0.32038 | 0.57316 | 0.65052 | 0.59388 | 0.75824 | 0.67996 | 0.53947 | 0.75591 | 0.16752 | 0.26803 | 0.12153 | 66.6667 | 17 | -21 | -1.1765 | 0 | -2 | 2.3204 | 2 | global_median | 0
F | 3 | F3 | 1 | U |  | 1 | 0.45346 | 0.33446 | 0.59244 | 0.63435 | 0.59228 | 0.73737 | 0.71484 | 0.5647 | 0.80345 | 0.14579 | 0.24818 | 0.095041 | 67.451 | 18 | -21 | -0.39216 | 1 | -2 | 2.2702 | 2.2361 | global_median | 0
F | 4 | F4 | 1 | U |  | 1 | 0.45908 | 0.32973 | 0.59244 | 0.61919 | 0.59024 | 0.71513 | 0.74142 | 0.55864 | 0.82844 | 0.12994 | 0.25287 | 0.08174 | 67.451 | 18 | -21 | -0.39216 | 1 | -2 | 2.2702 | 2.2361 | global_median | 0
F | 5 | F5 | 1 | U |  | 1 | 0.44787 | 0.32973 | 0.57316 | 0.60502 | 0.58778 | 0.6916 | 0.74026 | 0.56098 | 0.82874 | 0.13062 | 0.25106 | 0.081581 | 67.0588 | 17 | -19 | -0.78432 | 0 | 0 | 0.78432 | 0 | global_median | 0
G | 1 | G1 | 1 | U |  | 1 | 0.43134 | 0.31576 | 0.52344 | 0.64074 | 0.56622 | 0.73085 | 0.67319 | 0.55767 | 0.71621 | 0.17186 | 0.25363 | 0.14496 | 65.8824 | 15 | -17 | -1.9608 | -2 | 2 | 3.4416 | 2.8284 | global_median | 0
G | 2 | G2 | 1 | U |  | 1 | 0.42051 | 0.30663 | 0.54799 | 0.62174 | 0.56308 | 0.71286 | 0.67634 | 0.54457 | 0.76872 | 0.16984 | 0.26395 | 0.11423 | 65.098 | 17 | -20 | -2.7451 | 0 | -1 | 2.9216 | 1 | global_median | 0
G | 3 | G3 | 1 | U |  | 1 | 0.43134 | 0.31576 | 0.5605 | 0.60384 | 0.55952 | 0.69345 | 0.71433 | 0.56434 | 0.80827 | 0.1461 | 0.24846 | 0.092442 | 65.8824 | 17 | -20 | -1.9608 | 0 | -1 | 2.2011 | 1 | global_median | 0
G | 4 | G4 | 1 | U |  | 1 | 0.43134 | 0.31118 | 0.56681 | 0.58698 | 0.55555 | 0.6727 | 0.73485 | 0.56013 | 0.84259 | 0.1338 | 0.25171 | 0.074384 | 65.8824 | 17 | -21 | -1.9608 | 0 | -2 | 2.8008 | 2 | global_median | 0
G | 5 | G5 | 1 | U |  | 1 | 0.43681 | 0.32504 | 0.55423 | 0.57114 | 0.55116 | 0.65069 | 0.76481 | 0.58973 | 0.85176 | 0.11645 | 0.22935 | 0.069683 | 66.2745 | 16 | -19 | -1.5686 | -1 | 0 | 1.8603 | 1 | global_median | 0
H | 1 | H1 | 1 | U |  | 1 | 0.41515 | 0.31118 | 0.50543 | 0.60225 | 0.52443 | 0.66897 | 0.68933 | 0.59337 | 0.75554 | 0.16157 | 0.22667 | 0.12174 | 65.098 | 15 | -16 | -2.7451 | -2 | 3 | 4.5316 | 3.6056 | global_median | 0
H | 2 | H2 | 1 | U |  | 1 | 0.39929 | 0.29765 | 0.50543 | 0.58173 | 0.51943 | 0.65251 | 0.68638 | 0.57303 | 0.7746 | 0.16343 | 0.24182 | 0.11092 | 63.9216 | 16 | -17 | -3.9216 | -1 | 2 | 4.5143 | 2.2361 | global_median | 0
H | 3 | H3 | 1 | U |  | 1 | 0.40983 | 0.30213 | 0.5174 | 0.56235 | 0.51405 | 0.63468 | 0.72877 | 0.58774 | 0.81522 | 0.13741 | 0.23081 | 0.088725 | 64.3137 | 16 | -18 | -3.5294 | -1 | 1 | 3.8022 | 1.4142 | global_median | 0
H | 4 | H4 | 1 | U |  | 1 | 0.41515 | 0.30663 | 0.52344 | 0.54405 | 0.50827 | 0.61555 | 0.76306 | 0.60329 | 0.85036 | 0.11744 | 0.21947 | 0.070396 | 65.098 | 16 | -18 | -2.7451 | -1 | 1 | 3.088 | 1.4142 | global_median | 0
H | 5 | H5 | 1 | U |  | 1 | 0.41515 | 0.31576 | 0.52344 | 0.52681 | 0.50211 | 0.59521 | 0.78805 | 0.62887 | 0.87942 | 0.10345 | 0.20144 | 0.055804 | 65.4902 | 15 | -17 | -2.3529 | -2 | 2 | 3.6792 | 2.8284 | global_median | 0

## 05_REPLICATES_MEAN
ID | DF | Type | Conc | PAbs_Red_median | PAbs_Red_sd | PAbs_Green_median | PAbs_Green_sd | PAbs_Blue_median | PAbs_Blue_sd | L_median | L_sd | a_median | a_sd | b_median | b_sd | DeltaL_median | DeltaL_sd | Deltaa_median | Deltaa_sd | Deltab_median | Deltab_sd | DeltaE_ab_median | DeltaE_ab_sd | DeltaE_ab_chroma_median | DeltaE_ab_chroma_sd | CIELAB_ref_source | NReplicates | QCFlagged | QCCritical
1 | 1 | U |  | 0.13431 | 0.017247 | 0.23008 | 0.027961 | 0.09022 | 0.03013 | 67.8431 | 1.7442 | 17 | 1.4826 | -19 | 2.9652 | 0 | 1.7442 | 0 | 1.4826 | 0 | 2.9652 | 2.3204 | 1.3039 | 2 | 1.2282 | global_median | 40 | 1 | 0

## 06_FITTING
Channel | FitType | n_points | m | q | R2 | RMSE | ID | DF | C0 | C0_sd | beta_k | bias_index_k
PAbs_Red | Calibration | 0
PAbs_Red | UnknownFromEpsilon | 40 |  |  |  |  | 1 | 1 | 1.2304 | 0.158
PAbs_Green | Calibration | 0
PAbs_Green | UnknownFromEpsilon | 40 |  |  |  |  | 1 | 1 | 2.1077 | 0.25614
PAbs_Blue | Calibration | 0
PAbs_Blue | UnknownFromEpsilon | 40 |  |  |  |  | 1 | 1 | 0.82649 | 0.27602
L | Calibration | 0
L | UnknownOnly | 40 |  |  |  |  | 1 | 1
a | Calibration | 0
a | UnknownOnly | 40 |  |  |  |  | 1 | 1
b | Calibration | 0
b | UnknownOnly | 40 |  |  |  |  | 1 | 1
DeltaL | Calibration | 0
DeltaL | UnknownOnly | 40 |  |  |  |  | 1 | 1
Deltaa | Calibration | 0
Deltaa | UnknownOnly | 40 |  |  |  |  | 1 | 1
Deltab | Calibration | 0
Deltab | UnknownOnly | 40 |  |  |  |  | 1 | 1
DeltaE | Calibration | 0
DeltaE | UnknownOnly | 40 |  |  |  |  | 1 | 1
DeltaE_chroma | Calibration | 0
DeltaE_chroma | UnknownOnly | 40 |  |  |  |  | 1 | 1

## 07_METHOD_COMPARISON
Selected | Rank | Method | Family | ComparableGroup | CommonFactorsN | Score | ScoreFormula | RankMode | n_stdadd | n_unknown | Estimate_value | Estimate_sd | Estimate_source | expected_label_CRM | expected_id_CRM | expected_value_CRM | expected_sd_CRM | estimate_for_expected_CRM | delta_expected_CRM | recovery_pct_CRM | rel_error_CRM | expected_label_Colorimetry | expected_id_Colorimetry | expected_value_Colorimetry | expected_sd_Colorimetry | estimate_for_expected_Colorimetry | delta_expected_Colorimetry | recovery_pct_Colorimetry | rel_error_Colorimetry | BaseScore
1 | 1 | L | CIELAB | not_ranked | 0 | 0 | not_ranked |  | 0 | 0 |  |  |  | CRM | 1 | 1.92 | 0.01 |  |  |  |  | Colorimetry | 1 | 2.1 | 0.15 |  |  |  |  | 0
0 | 2 | a | CIELAB | not_ranked | 0 | 0 | not_ranked |  | 0 | 0 |  |  |  | CRM | 1 | 1.92 | 0.01 |  |  |  |  | Colorimetry | 1 | 2.1 | 0.15 |  |  |  |  | 0
0 | 3 | b | CIELAB | not_ranked | 0 | 0 | not_ranked |  | 0 | 0 |  |  |  | CRM | 1 | 1.92 | 0.01 |  |  |  |  | Colorimetry | 1 | 2.1 | 0.15 |  |  |  |  | 0
0 | 4 | DeltaE | DeltaCIELAB | not_ranked | 0 | 0 | not_ranked |  | 0 | 0 |  |  |  | CRM | 1 | 1.92 | 0.01 |  |  |  |  | Colorimetry | 1 | 2.1 | 0.15 |  |  |  |  | 0
0 | 5 | DeltaE_chroma | DeltaCIELAB | not_ranked | 0 | 0 | not_ranked |  | 0 | 0 |  |  |  | CRM | 1 | 1.92 | 0.01 |  |  |  |  | Colorimetry | 1 | 2.1 | 0.15 |  |  |  |  | 0
0 | 6 | DeltaL | DeltaCIELAB | not_ranked | 0 | 0 | not_ranked |  | 0 | 0 |  |  |  | CRM | 1 | 1.92 | 0.01 |  |  |  |  | Colorimetry | 1 | 2.1 | 0.15 |  |  |  |  | 0
0 | 7 | Deltaa | DeltaCIELAB | not_ranked | 0 | 0 | not_ranked |  | 0 | 0 |  |  |  | CRM | 1 | 1.92 | 0.01 |  |  |  |  | Colorimetry | 1 | 2.1 | 0.15 |  |  |  |  | 0
0 | 8 | Deltab | DeltaCIELAB | not_ranked | 0 | 0 | not_ranked |  | 0 | 0 |  |  |  | CRM | 1 | 1.92 | 0.01 |  |  |  |  | Colorimetry | 1 | 2.1 | 0.15 |  |  |  |  | 0
0 | 9 | PAbs_Blue | RGB | not_ranked | 0 | 0 | not_ranked | unavailable | 0 | 1 | 0.82649 | 0.27602 | unknown_from_calibration | CRM | 1 | 1.92 | 0.01 | 0.82649 | -1.0935 | 43.0465 | -0.56953 | Colorimetry | 1 | 2.1 | 0.15 | 0.82649 | -1.2735 | 39.3568 | -0.60643 | 0
0 | 10 | PAbs_Green | RGB | not_ranked | 0 | 0 | not_ranked | unavailable | 0 | 1 | 2.1077 | 0.25614 | unknown_from_calibration | CRM | 1 | 1.92 | 0.01 | 2.1077 | 0.18773 | 109.7777 | 0.097777 | Colorimetry | 1 | 2.1 | 0.15 | 2.1077 | 0.0077309 | 100.3681 | 0.0036814 | 0
0 | 11 | PAbs_Red | RGB | not_ranked | 0 | 0 | not_ranked | unavailable | 0 | 1 | 1.2304 | 0.158 | unknown_from_calibration | CRM | 1 | 1.92 | 0.01 | 1.2304 | -0.68958 | 64.0842 | -0.35916 | Colorimetry | 1 | 2.1 | 0.15 | 1.2304 | -0.86958 | 58.5913 | -0.41409 | 0

## 08_LEGENDS
Term | Meaning | Formula | Unit | Where used | Shown when | Notes
a | CIELAB green–red axis | median ROI value after RGB → CIELAB conversion | dimensionless | RAW, SUMMARY | CIELAB enabled
b | CIELAB blue–yellow axis | median ROI value after RGB → CIELAB conversion | dimensionless | RAW, SUMMARY | CIELAB enabled
BaseScore/FinalScore | Intermediate/final score fields when present | same common score unless an explicit post-score adjustment is applied | dimensionless | METHOD_COMPARISON | method comparison | If identical to Score in the exported workbook, Score is the authoritative field.
beta_k | Slope ratio for one standard-addition curve | m_std / m_cal | dimensionless | FITTING, METHOD_COMPARISON | stdadd present | Used to assess whether StdAdd and calibration slopes are coherent.
beta_k/beta_mean | Standard-addition/calibration slope ratio | m_std / m_cal; beta_mean is the mean across available standard-addition fits | dimensionless | 06_FITTING, 07_METHOD_COMPARISON | calibration plus standard addition present
bias_index_k | Absolute relative slope bias for one standard-addition curve | ¦beta_k − 1¦ | dimensionless | FITTING, METHOD_COMPARISON | stdadd present | 0 means identical slopes.
bias_index_k/bias_index_mean | Relative slope-bias index | ¦m_std/m_cal - 1¦; bias_index_mean is the mean across available fits | dimensionless | 06_FITTING, 07_METHOD_COMPARISON | calibration plus standard addition present
BiasIndex | Relative slope bias between standard addition and calibration | ¦β − 1¦, where β = m_std / m_cal | dimensionless | METHOD_COMPARISON | stdadd present | Equals 0 when the standard-addition slope equals the calibration slope. Larger values indicate stronger slope bias.
C0 | Concentration in original sample | standard addition: DF × (q/m); unknown/CRM from calibration: DF × ((y − q)/m); unknown from epsilon: DF × PAbs/(epsilon × l_cm), converted from M to the selected unit | mM | FITTING, OVERVIEW | stdadd present or unknown/CRM present | Same symbol is used for standard addition and unknown/CRM results; FitType defines the method.
C0/C0_sd | Estimated original-sample concentration and associated uncertainty | standard addition: DF x q/m; calibration projection: DF x (y - q)/m; epsilon mode: DF x PAbs/(epsilon x l_cm), converted from M to the selected unit | mM | FIGURE_RGB.png, 03_OVERVIEW, 06_FITTING | standard addition, unknown or epsilon quantification present | FitType identifies which calculation path produced C0.
C0_sd | Uncertainty on C0 | Fit covariance for calibration/std-add; for epsilon mode, propagated PAbs replicate SD divided by epsilon × l_cm and converted to the selected unit | mM | FITTING, OVERVIEW | stdadd present or unknown/CRM present | Reported in the same concentration units as C0.
Channel | Fitted analytical descriptor | PAbs_Red, PAbs_Green, PAbs_Blue or diagnostic CIELAB/DeltaE descriptor | text | FITTING, METHOD_COMPARISON | always
CIELAB (L*, a*, b*) | CIE 1976 color coordinates derived from the RGB image | L* = 116 f(Y/Yn) - 16; a* = 500[f(X/Xn) - f(Y/Yn)]; b* = 200[f(Y/Yn) - f(Z/Zn)] | dimensionless | RAW, REPLICATES_MEAN, CIELAB diagnostics | CIELAB enabled | D65 is the reference white used by the sRGB/OpenCV conversion convention. References: CIE 1976 L*a*b* and CIE standard illuminant D65.
CIELAB_ref_source | Source of the CIELAB reference used to compute Delta variables | zero calibration or other reported fallback source | text | 04_RAW, 05_REPLICATES_MEAN | CIELAB/DeltaE outputs present | Reported only for CIELAB-derived variables because RGB pseudo-absorbance does not use this CIELAB reference.
ClipDelta | Positive calibration-derived clipping correction | max(0, y_expected − y_shifted) when deficit exceeds SD threshold | signal | FITTING | RGB calibration | Computed from calibration after S0 and zero-intercept fit; applied by same channel and same x.
ClipX | Calibration x values where clipping correction was evaluated | list of calibration concentrations | mM | FITTING | RGB calibration | Transferred to StdAdd at the same channel and same x values.
Code implementation | Main computational libraries and code provenance | custom Python code using NumPy, OpenCV, openpyxl and matplotlib | software provenance | all generated outputs | always | Sources/libraries: NumPy for least-squares solves; OpenCV for image masks/color conversion; openpyxl for xlsx export; matplotlib for PNG figures.
Col | Plate column index | 1-based column number | well-column index | RAW, REPLICATES_MEAN | always
CommonFactorsN | Number of common factors used in the score | 3 for R²_cal, R²_std, slope agreement; 1 for calibration-only or stdadd-only fallbacks | integer | METHOD_COMPARISON | always | Used to avoid comparing scores obtained from different formulas.
ComparableGroup | Set of methods scored with the same formula | calibration_plus_stdadd, calibration_only, stdadd_only, or not_ranked | text | METHOD_COMPARISON | always | Scores are directly comparable only within the same group. The group with the largest CommonFactorsN is used for Selected/Rank.
Conc | Configured concentration for calibration or added-standard wells | user/configurator input | mM | RAW, REPLICATES_MEAN, FITTING | always | Unknown wells may have undefined Conc.
Confidence class | Qualitative class derived from reliability score | HIGH >= 75; MEDIUM >= 45; LOW < 45; NOT QUANTIFIABLE when no valid quantification is available | class | OVERVIEW | always | Interpret together with the reliability reason.
delta_expected_*/delta_reference_* | Difference between estimate and external reference | estimate - reference | mM | METHOD_COMPARISON.png, 03_OVERVIEW, 07_METHOD_COMPARISON | reference values configured
delta_reference_* | Difference between estimated concentration and an external reference | C0 - reference_value | mM | OVERVIEW, METHOD_COMPARISON, FIGURE_RGB | reference values configured
Deltaa | Difference in CIELAB a* relative to the selected reference | a - a_ref | dimensionless | RAW, SUMMARY | CIELAB enabled | Reference source is reported in CIELAB_ref_source.
Deltab | Difference in CIELAB b* relative to the selected reference | b - b_ref | dimensionless | RAW, SUMMARY | CIELAB enabled | Reference source is reported in CIELAB_ref_source.
DeltaE_ab | Total CIELAB color difference | sqrt(DeltaL^2 + Deltaa^2 + Deltab^2) | dimensionless | FIGURE_CIELAB_DELTAE.png, RAW_DATA_DETAILS_CAPTION.txt, 04_RAW, 05_REPLICATES_MEAN, 07_METHOD_COMPARISON, diagnostics | DeltaE outputs present | Reference: CIE 1976 L*a*b* color-difference form.
DeltaE_ab_chroma | Chromatic CIELAB color difference without lightness | sqrt(Deltaa^2 + Deltab^2) | dimensionless | FIGURE_CIELAB_DELTAE.png, RAW_DATA_DETAILS_CAPTION.txt, 04_RAW, 05_REPLICATES_MEAN, diagnostics | DeltaE outputs present | Reference source is reported in CIELAB_ref_source.
DeltaL | Difference in CIELAB lightness relative to the selected reference | L - L_ref | dimensionless | RAW, SUMMARY | CIELAB enabled | Reference source is reported in CIELAB_ref_source.
DeltaL/Deltaa/Deltab | CIELAB coordinate differences from the selected CIELAB reference | DeltaL = L - L_ref; Deltaa = a - a_ref; Deltab = b - b_ref | dimensionless | FIGURE_CIELAB_DELTAE.png, 04_RAW, 05_REPLICATES_MEAN, diagnostics | Delta outputs present | The reference source is reported in CIELAB_ref_source.
DF | Dilution factor used to report original-sample concentration | C0 = DF x concentration in analyzed/diluted well | dimensionless | RAW, REPLICATES_MEAN, FITTING | always
Empty-well QC | QC based on empty-well robust SD and/or stored empty comparison | warning/watch/ok/not_available from drift score and robust SD thresholds | class | OVERVIEW | always | Flags comparability or background/empty-well drift.
epsilon | User-configured Beer-Lambert-like proportionality coefficient for PAbs quantification | PAbs = epsilon × l_cm × C_M | M-1 cm-1 | FITTING, OVERVIEW | unknown-only epsilon mode | C_M is mol/L. The analyzer converts the calculated concentration to the selected output unit.
Estimate_value/Estimate_sd | Representative concentration estimate and associated SD used for method comparison | derived from standard addition, calibration projection, or epsilon mode depending on Estimate_source | mM | METHOD_COMPARISON | method comparison
Estimate_value/Estimate_sd/Estimate_source | Representative concentration estimate, associated SD and source | selected according to Estimate_source | mM and text | METHOD_COMPARISON.png, 07_METHOD_COMPARISON | method comparison present
expected_*/reference_* | External reference metadata configured by the user | user/configurator input | text or mM | METHOD_COMPARISON.png, 03_OVERVIEW, 07_METHOD_COMPARISON | reference values configured | External reference values are checks only and are not used for ranking.
Family | Method family used in method comparison | RGB, CIELAB, DeltaCIELAB, or other | text | METHOD_COMPARISON | always
Field | Name of a metadata or overview item | label reported in the first column of the sheet | text | 02_METADATA, 03_OVERVIEW | always
FitType | Type of fit or concentration-estimation row | Calibration, StdAdd, UnknownFromCal, UnknownFromEpsilon, UnknownOnly | text | FITTING | always
flatfield_span | Image-level slow-field nonuniformity | (P95 − P5) / median of the slow grayscale field | dimensionless | METADATA | always | Part of rule-based image QC; larger values indicate stronger illumination/background gradient.
FracWellWarnings/FracWellCritical | Fractions of replicate wells with warning or critical optical QC | count / NReplicates | dimensionless | REPLICATES_MEAN | always
Geometry and epsilon/path-length quantification | Assumption used when epsilon-based unknown quantification is enabled | l_cm = (volume_uL / well_bottom_area_mm2) / 10; C_M = PAbs / (epsilon x l_cm) | cm and M | RESULTS_CAPTION.txt, RAW_DATA_DETAILS_CAPTION.txt, REPORT | epsilon/path-length mode configured | Assumes ANSI/SLAS-compatible flat-bottom microplate geometry; non-flat or non-certified geometries require separate validation.
ID | Sample or reference identifier assigned in the plate map | user/configurator input | text | RAW, REPLICATES_MEAN, FITTING, METHOD_COMPARISON | always
ImageWarning | Well-level optical QC warning flag | 1 if optical QC rules flag the well, else 0 | 0/1 | 04_RAW | always
ImageWarning_any | Group-level indicator that at least one replicate had an image warning | 1 if any replicate ImageWarning = 1 | 0/1 | REPLICATES_MEAN | always
IRLS | Iteratively reweighted least-squares robust linear regression with residual-based weights | minimize sum_i w_i (y_i - (m x_i + q))^2; w_i is updated iteratively from residual magnitude | dimensionless | FIGURE_RGB.png, FIGURE_CIELAB_DELTAE.png, RESULTS_CAPTION.txt, RAW_DATA_DETAILS_CAPTION.txt, 06_FITTING, 07_METHOD_COMPARISON, diagnostic fitting tables | always | Reference: Huber, P. J. (1964), Robust Estimation of a Location Parameter. Implementation: custom NumPy-based IRLS using repeated least-squares solves.
L | CIELAB lightness | median ROI value after RGB → CIELAB conversion | dimensionless | RAW, SUMMARY | CIELAB enabled
L/a/b | CIELAB lightness and opponent-color coordinates | median ROI value after RGB to CIELAB conversion | dimensionless | FIGURE_CIELAB_DELTAE.png, 04_RAW, 05_REPLICATES_MEAN, diagnostics | CIELAB outputs present
liquid_volume_ul | Liquid volume loaded into each well | user input | uL | OVERVIEW, FITTING | epsilon mode configured | Used only to calculate path length for flat-bottom wells.
LOD | Limit of detection | 3 × σ / ¦m¦ | mM | FITTING, OVERVIEW | calibration only or calibration present | Shown only for calibration rows.
LOD/LOQ | Detection and quantification limits from calibration | LOD = 3 x sigma_cal / ¦m¦; LOQ = 10 x sigma_cal / ¦m¦ | mM | FIGURE_RGB.png, FIGURE_CIELAB_DELTAE.png, 03_OVERVIEW, 06_FITTING, 07_METHOD_COMPARISON | calibration present
LOQ | Limit of quantification | 10 × σ / ¦m¦ | mM | FITTING, OVERVIEW | calibration only or calibration present | Shown only for calibration rows.
m | Slope of linear fit | y = m x + q | signal / mM | FITTING | always | Calibration and standard-addition fits use the same notation.
m/q | Slope and intercept of the linear fit | y = m x + q | response / concentration and response | FIGURE_RGB.png, FIGURE_CIELAB_DELTAE.png, 06_FITTING | fit rows present
MeanBG_* | Linearized local background intensity for the channel | background surface evaluated at the well and gamma linearized | dimensionless | RAW | always | Derived from inter-well background reconstruction.
MeanBG_Red/MeanBG_Green/MeanBG_Blue | Linearized local inter-well background intensity for each RGB channel | 2D background surface evaluated at the well and gamma-linearized | dimensionless | 04_RAW | always
MeanW_* | Linearized median well intensity for the channel | median ROI channel intensity after gamma linearization | dimensionless | RAW | always | Computed from selected well ROI pixels.
MeanW_Red/MeanW_Green/MeanW_Blue | Linearized median well intensity for each RGB channel | median ROI channel intensity after gamma linearization | dimensionless | 04_RAW | always
Method/Family/ComparableGroup/RankMode | Method-comparison identifiers and comparable-score grouping | text labels defining descriptor, family and score comparability | text | METHOD_COMPARISON.png, 07_METHOD_COMPARISON | method comparison present | Scores should be compared directly only within the same ComparableGroup.
n_points | Number of finite data points used by the fit | count of finite x,y pairs | count | FITTING | always | All finite points are retained by robust IRLS.
NClipPoints/ClipX/ClipDelta | Calibration clipping/baseline diagnostic fields | NClipPoints = number of adjusted points; ClipX = concentrations; ClipDelta = applied response adjustments | count, concentration, response | 06_FITTING | calibration present | Diagnostic only; clipping is not part of the common method score.
NReplicates | Number of replicate wells in a summarized group | count of wells in group | count | 05_REPLICATES_MEAN | always
NWellWarnings/NWellCritical | Counts of replicate wells with warning or critical optical QC | count over replicate wells | count | REPLICATES_MEAN | always
PAbs | Image-derived RGB pseudo-absorbance | PAbs = log10(I_BG / I_well) = -log10(I_well / I_BG) | dimensionless | FIGURE_RGB.png, RESULTS_CAPTION.txt, 04_RAW, 05_REPLICATES_MEAN, 06_FITTING, 07_METHOD_COMPARISON | always | Pseudo-absorbance is not assumed to be spectrophotometric absorbance.
PAbs_Blue | RGB pseudo-absorbance (blue channel) | log10(BG_blue/W_blue) = -log10(W_blue/BG_blue) | dimensionless | RAW, REPLICATES_MEAN, FITTING | always | Image-derived pseudo-absorbance with fixed full-background normalization.
PAbs_Green | RGB pseudo-absorbance (green channel) | log10(BG_green/W_green) = -log10(W_green/BG_green) | dimensionless | RAW, REPLICATES_MEAN, FITTING | always | Image-derived pseudo-absorbance with fixed full-background normalization.
PAbs_Red | RGB pseudo-absorbance (red channel) | log10(BG_red/W_red) = -log10(W_red/BG_red) | dimensionless | RAW, REPLICATES_MEAN, FITTING | always | Image-derived pseudo-absorbance with fixed full-background normalization.
PAbs_Red/PAbs_Green/PAbs_Blue | RGB pseudo-absorbance for the red, green and blue channels | log10(MeanBG_channel / MeanW_channel) | dimensionless | FIGURE_RGB.png, 04_RAW, 05_REPLICATES_MEAN, 06_FITTING, 07_METHOD_COMPARISON | always | Exported in standard RGB order even though OpenCV stores images internally as BGR.
path_length | Estimated optical path length | l_cm = (V_uL / A_bottom_mm2) / 10 | cm | OVERVIEW, FITTING | epsilon mode configured | Assumes 1 uL = 1 mm^3 and a flat-bottom well with nominal bottom area.
path_length_mm | Estimated liquid height | l_mm = V_uL / A_bottom_mm2 | mm | OVERVIEW, FITTING | epsilon mode configured | Converted to cm before Beer-Lambert-like concentration calculation.
Purpose | Short description of the sheet role | free-text description | text | 01_CONTENTS | always
q | Intercept of linear fit | y = m x + q | signal | FITTING | always | For standard addition, C0 is derived from m and q.
QCCritical | Group-level critical optical quality flag | 1 if FracWellCritical >= 0.50 | 0/1 | REPLICATES_MEAN | always | Critical flag; points remain reported and used.
QCFlagged | Group-level optical quality warning | 1 if any replicate has ImageWarning, any critical replicate, or median UsedFraction < 0.70 | 0/1 | REPLICATES_MEAN | always | Warning flag; points remain reported and used.
QCFlagged/QCCritical | Replicate-group QC warning and critical flags | rule-based aggregation of well-level warnings, critical wells and UsedFraction | 0/1 | 05_REPLICATES_MEAN | always
R2 | Coefficient of determination | 1 - SSE/SST | dimensionless | FIGURE_RGB.png, FIGURE_CIELAB_DELTAE.png, 06_FITTING, 07_METHOD_COMPARISON | fit rows present | Closer to 1 indicates better fit.
R2_cal/R2_std_mean/m_cal/m_std_mean | Calibration and standard-addition fit descriptors used in method comparison | R2_cal from calibration fit; R2_std_mean = mean standard-addition R2; m_cal and m_std_mean are fit slopes | mixed | METHOD_COMPARISON.png, 07_METHOD_COMPARISON | method comparison present
Rank | Workbook-level rank of each method | 1 = highest Score within the selected ComparableGroup | integer | METHOD_COMPARISON | always | Rows outside the selected ComparableGroup are not assigned a direct rank.
RankMode | Data basis available for ranking | calibration_plus_stdadd, calibration_only, stdadd_only, or unavailable | text | METHOD_COMPARISON | always
recovery_pct_* | Recovery relative to external reference | 100 x estimate / reference | % | FIGURE_RGB.png, METHOD_COMPARISON.png, 03_OVERVIEW, 07_METHOD_COMPARISON | reference values configured
rel_error_* | Relative error versus external reference | 100 x (estimate - reference) / reference | % | 03_OVERVIEW, 07_METHOD_COMPARISON | reference values configured
Reliability criteria | Rule used to flag a method as low reliability in method comparison | low if slope agreement < 0.5 OR R² < 0.8 OR ¦Δ¦ > threshold | rule | METHOD_COMPARISON | stdadd present | The Δ threshold is max(3×reference SD, 50% of the reference value), evaluated against the nearest available reference.
Reliability score | Overall report-level reliability score | starts at 50 and adds/subtracts rule-based penalties/bonuses for calibration availability, stored calibration, plate QC, critical wells, used fraction, empty-well QC, and ranking separation | 0-100 | OVERVIEW | always | Not a statistical probability; it is a heuristic audit score.
RGB linearization | Conversion of gamma-domain RGB values to an intensity-like scale | I_lin = (I/255)^gamma, gamma approximately 2.2 | dimensionless | PAbs calculation in 04_RAW and FIGURE_RGB.png | always | No automatic image correction is applied.
RGB to CIELAB | Diagnostic conversion from image RGB/BGR values to CIE L*a*b* descriptors | OpenCV BGR->Lab for 8-bit sRGB-like images; reference form: linearized sRGB -> XYZ using the standard sRGB-to-XYZ matrix and D65 reference white -> CIE L*a*b* | dimensionless | FIGURE_CIELAB_DELTAE.png, RAW_DATA_DETAILS_CAPTION.txt, 04_RAW, 05_REPLICATES_MEAN, DIAGNOSTICS | CIELAB/DeltaE outputs present | References: OpenCV color conversions; IEC 61966-2-1:1999 sRGB; CIE 1976 L*a*b*; CIE standard illuminant D65.
RMSE | Root mean squared error of the fit | sqrt(mean(residual^2)) | response unit | 06_FITTING | fit rows present | Interpret in the units of the fitted response.
Robust SD | Robust dispersion across replicates | 1.4826 × MAD, where MAD = median(¦x − median(x)¦) | same as variable | REPLICATES_MEAN | always | Computed within each replicate group.
Row | Plate row label | A-based alphabetical row label | well-row label | RAW, REPLICATES_MEAN | always
S0_calibration | Channel-wide low-signal offset estimated from raw calibration | max(0, −min(raw calibration response)) | signal | FITTING | RGB calibration | Estimated before forced-zero calibration and transferred to StdAdd/unknown data.
Score | Common method-ranking score | for calibration plus standard addition: SlopeAgreement^2 x sqrt(R2_cal x R2_std) | dimensionless | FIGURE_RGB.png, METHOD_COMPARISON.png, 07_METHOD_COMPARISON | method comparison present | Expected/reference values, recovery, SNR and clipping are not used in this score.
ScoreFormula | Formula used to compute Score | text descriptor | text | METHOD_COMPARISON | always | Documents which score formula was applied to the row.
Selected | Workbook-level selected method flag | 1 for the highest-ranked row within the most informative ComparableGroup, otherwise 0 | 0/1 | METHOD_COMPARISON | always | Selection in the workbook is derived from method-comparison rank and not from expected/recovery.
Selected/Rank | Selected method flag and rank position | Selected = 1 for the selected method; Rank = score order within the highest common-factor group | 0/1 and integer | 07_METHOD_COMPARISON | method comparison present
Sheet | Workbook sheet name | worksheet name | text | 01_CONTENTS | always
sigma_cal | Noise estimate for calibration | SD at zero concentration or median calibration SD | signal | FITTING | calibration only or calibration present | Source is reported in sigma_source.
sigma_cal/sigma_source/SNR | Calibration noise estimate, its source and slope-to-noise ratio | SNR = ¦m¦ / sigma_cal | mixed | 06_FITTING, 07_METHOD_COMPARISON | calibration present | SNR is diagnostic only and is not used in the common score.
sigma_source | Source of sigma_cal | median_calibration_sd, blank_zero_calibration_sd, or unavailable | text | FITTING | calibration present
SignalT_* | Transmittance-like intensity ratio before logarithm | MeanW_* / MeanBG_* | dimensionless | RAW | always | PAbs_* = -log10(SignalT_*).
SignalT_Red/SignalT_Green/SignalT_Blue | Transmittance-like ratio before logarithmic conversion | SignalT_channel = MeanW_channel / MeanBG_channel | dimensionless | 04_RAW | always | PAbs_channel = -log10(SignalT_channel).
SlopeAgreement | Agreement between calibration and standard-addition slope magnitudes | min(¦m_cal¦, ¦m_std¦) / max(¦m_cal¦, ¦m_std¦) | dimensionless | METHOD_COMPARISON.png, 07_METHOD_COMPARISON | calibration plus standard addition present | Equals 1 when slopes are identical in magnitude; lower values indicate poorer agreement between calibration and standard addition.
SNR | Slope-to-noise ratio | ¦m¦ / sigma_cal | dimensionless | FITTING | calibration only or calibration present | Used only for calibration rows.
Type | Well role in the analytical workflow | C = calibration; A = standard addition; U = unknown; other labels may be reported when present | text | RAW, REPLICATES_MEAN | always
UsedFraction | Fraction of ROI core pixels retained after filtering | n_used / n_core | dimensionless | RAW, SUMMARY | always | Lower values indicate stronger filtering. Closer to 1 means less aggressive trimming.
UsedFraction_median | Median UsedFraction across replicate wells in a group | median(UsedFraction) | dimensionless | REPLICATES_MEAN | always
Value | Value associated with a metadata or overview item | reported value for the corresponding Field | mixed | 02_METADATA, 03_OVERVIEW | always
Well | Human-readable well identifier | Well = Row + Col | well label | RAW, REPLICATES_MEAN | always
well_bottom_area_mm2 | Nominal flat-bottom well area used for path-length calculation | pi × (d_bottom/2)^2 | mm^2 | OVERVIEW | epsilon mode configured | Derived from the selected plate geometry.