## 01_CONTENTS
Sheet | Purpose
01_CONTENTS | Index of primary results workbook sheets.
02_METADATA | Image-level metadata and rule-based image QC used to audit the analysis.
03_OVERVIEW | Final quantitative summary, selected method, reliability and external reference checks when provided.
04_RAW | Well-level analytical values used by the fitting pipeline.
05_REPLICATES_MEAN | Replicate-group medians, robust SDs and group QC flags.
06_FITTING | Calibration, standard-addition and unknown/CRM fit results.
07_METHOD_COMPARISON | Method ranking using only common score factors; expected values are external checks only.
08_LEGENDS | Definitions for all fields reported in this workbook and primary RGB figure.

## 02_METADATA
Field | Value | Notes
Analysis image source | original | original or in-memory resized image used for all measurements
Original image path | C:/Users/prova/OneDrive - unifi.it/Desktop/cromo/Piastra viola/U2.jpg | source file selected by the user
Original image size | 1606 x 1094 px | width x height in pixels, when known
Analysis image size | 1606 x 1094 px | width x height in pixels actually analyzed
Resize scale | 1 | analysis_width / original_width, when a resized image is used
Initial image QC | WARNING | Rule-based result: FAIL if destructive=True; WARNING if quality_warning=True or borderline=True; otherwise OK. No image correction is applied.
Initial image QC class | quality_warning | non_correctable if destructive=True; quality_warning if flatfield_span > 0.18 and no destructive defect; usable_with_warnings if only borderline defects; good if no flags.
Initial image QC messages | slow illumination/background gradient; borderline blur | Messages list the threshold rules triggered by the measured fields below.
Image QC decision rules | thresholds | destructive = dead_channel OR saturation_fraction>0.003 OR saturation_all_channels_fraction>0.0008 OR specular_fraction>0.003 OR max_side<900 OR approx_roi_pixels<120; borderline = saturation_fraction>0.0005 OR specular_fraction>0.0005 OR blur_score<35; quality_warning = flatfield_span>0.18.
Flat-field span | 0.36633 | Flatfield_span = (P95 - P5) / median of the slow illumination field estimated from the grayscale image. Larger values indicate stronger illumination/background gradient.
Specular fraction | 8.822e-05 | fraction of bright low-saturation pixels
Saturation fraction | 8.0821e-05 | fraction of pixels with at least one saturated channel
Dead channel | 0 | 1 if a channel is near-dead by mean/variance checks
Approx well pitch | 112.4306 | median distance between adjacent well centers, in pixels
Approx ROI pixels/well | 3526 | median number of pixels used in the final floor ROI mask per well
Blur score | 10.0098 | variance of Laplacian of the analysis image; lower values indicate stronger blur

## 03_OVERVIEW
Field | Value
Quantification | available
Reliability score | 17
Confidence class | LOW
Reliability note | concentration estimated from PAbs with epsilon because no calibration was available; no quantitative method was ranked as available; epsilon-based quantification without calibration has intrinsically limited reliability
Empty-well QC status | ok
Configured epsilon (M-1 cm-1) | 180
Liquid volume per well (uL) | 200
Calculated path length (cm) | 0.60644
Calculated liquid height (mm) | 6.0644
Path-length source | calculated_from_volume_and_flat_bottom_geometry
Plate geometry for path length | 96-well standard F-bottom / flat-bottom
Plate geometry assumption | ANSI/SLAS-compatible flat-bottom geometry; vendor dimensions may vary by plate family
Selected method from rank | L
Selected family | CIELAB
Ranking mode
Selected method score | 0
Status notes | epsilon quantification used; empty-well QC: ok; epsilon configured; path length calculated_from_volume_and_flat_bottom_geometry

## 04_RAW
Row | Col | Well | ID | Type | Conc | DF | MeanW_Red | MeanW_Green | MeanW_Blue | MeanBG_Red | MeanBG_Green | MeanBG_Blue | SignalT_Red | SignalT_Green | SignalT_Blue | PAbs_Red | PAbs_Green | PAbs_Blue | L | a | b | DeltaL | Deltaa | Deltab | DeltaE_ab | DeltaE_ab_chroma | CIELAB_ref_source | ImageWarning
A | 1 | A1 | 1 | U |  | 1 | 0.32504 | 0.2428 | 0.35374 | 0.37003 | 0.27516 | 0.42825 | 0.87842 | 0.88238 | 0.82601 | 0.0563 | 0.054346 | 0.083013 | 58.4314 | 12 | -10 | 2.9412 | -3 | 3 | 5.1624 | 4.2426 | global_median | 0
A | 2 | A2 | 1 | U |  | 1 | 0.30663 | 0.22331 | 0.33922 | 0.39913 | 0.31163 | 0.44637 | 0.76826 | 0.71658 | 0.75995 | 0.11449 | 0.14473 | 0.11921 | 56.4706 | 13 | -12 | 0.98039 | -2 | 1 | 2.4416 | 2.2361 | global_median | 0
A | 3 | A3 | 1 | U |  | 1 | 0.32504 | 0.22331 | 0.35374 | 0.42789 | 0.34961 | 0.46413 | 0.75963 | 0.63874 | 0.76215 | 0.1194 | 0.19468 | 0.11796 | 57.2549 | 14 | -12 | 1.7647 | -1 | 1 | 2.2615 | 1.4142 | global_median | 0
A | 4 | A4 | 1 | U |  | 1 | 0.35374 | 0.2428 | 0.38377 | 0.45614 | 0.38894 | 0.48148 | 0.77551 | 0.62424 | 0.79707 | 0.11041 | 0.20465 | 0.098502 | 59.2157 | 15 | -12 | 3.7255 | 0 | 1 | 3.8574 | 1 | global_median | 0
A | 5 | A5 | 1 | U |  | 1 | 0.37868 | 0.27158 | 0.39929 | 0.48372 | 0.42948 | 0.49837 | 0.78284 | 0.63234 | 0.8012 | 0.10632 | 0.19905 | 0.09626 | 61.5686 | 14 | -11 | 6.0784 | -1 | 2 | 6.4767 | 2.2361 | global_median | 0
B | 1 | B1 | 1 | U |  | 1 | 0.34403 | 0.23883 | 0.3636 | 0.36273 | 0.2684 | 0.41895 | 0.94843 | 0.88982 | 0.8679 | 0.022996 | 0.050699 | 0.06153 | 58.4314 | 15 | -12 | 2.9412 | 0 | 1 | 3.1065 | 1 | global_median | 0
B | 2 | B2 | 1 | U |  | 1 | 0.30213 | 0.20471 | 0.33446 | 0.39141 | 0.30346 | 0.43716 | 0.77189 | 0.67459 | 0.76507 | 0.11244 | 0.17096 | 0.1163 | 55.2941 | 15 | -13 | -0.19608 | 0 | 0 | 0.19608 | 0 | global_median | 0
B | 3 | B3 | 1 | U |  | 1 | 0.31576 | 0.20471 | 0.34886 | 0.41974 | 0.33991 | 0.45503 | 0.75228 | 0.60225 | 0.76668 | 0.12362 | 0.22022 | 0.11539 | 55.6863 | 16 | -14 | 0.19608 | 1 | -1 | 1.4277 | 1.4142 | global_median | 0
B | 4 | B4 | 1 | U |  | 1 | 0.34886 | 0.22714 | 0.38377 | 0.44756 | 0.3776 | 0.47251 | 0.77948 | 0.60153 | 0.8122 | 0.1082 | 0.22074 | 0.090336 | 58.0392 | 17 | -14 | 2.549 | 2 | -1 | 3.3908 | 2.2361 | global_median | 0
B | 5 | B5 | 1 | U |  | 1 | 0.37868 | 0.25903 | 0.39929 | 0.47471 | 0.41637 | 0.48955 | 0.79769 | 0.6221 | 0.81563 | 0.098165 | 0.20614 | 0.088507 | 60.7843 | 15 | -12 | 5.2941 | 0 | 1 | 5.3877 | 1 | global_median | 0
C | 1 | C1 | 1 | U |  | 1 | 0.32973 | 0.22331 | 0.34886 | 0.35451 | 0.26317 | 0.40816 | 0.9301 | 0.84853 | 0.85473 | 0.031472 | 0.071334 | 0.068169 | 57.2549 | 15 | -12 | 1.7647 | 0 | 1 | 2.0283 | 1 | global_median | 0
C | 2 | C2 | 1 | U |  | 1 | 0.29322 | 0.19397 | 0.32504 | 0.38272 | 0.29695 | 0.42642 | 0.76614 | 0.65322 | 0.76225 | 0.11569 | 0.18494 | 0.1179 | 54.1176 | 15 | -14 | -1.3725 | 0 | -1 | 1.6982 | 1 | global_median | 0
C | 3 | C3 | 1 | U |  | 1 | 0.29765 | 0.19397 | 0.33446 | 0.41058 | 0.332 | 0.44435 | 0.72496 | 0.58426 | 0.75269 | 0.13969 | 0.23339 | 0.12339 | 54.1176 | 17 | -14 | -1.3725 | 2 | -1 | 2.6237 | 2.2361 | global_median | 0
C | 4 | C4 | 1 | U |  | 1 | 0.32504 | 0.20836 | 0.35865 | 0.43794 | 0.36817 | 0.46192 | 0.7422 | 0.56594 | 0.77644 | 0.12948 | 0.24723 | 0.10989 | 56.0784 | 17 | -15 | 0.58824 | 2 | -2 | 2.8889 | 2.8284 | global_median | 0
C | 5 | C5 | 1 | U |  | 1 | 0.35374 | 0.23883 | 0.38377 | 0.46462 | 0.40532 | 0.47906 | 0.76135 | 0.58924 | 0.80109 | 0.11842 | 0.22971 | 0.096318 | 58.8235 | 15 | -13 | 3.3333 | 0 | 0 | 3.3333 | 0 | global_median | 0
D | 1 | D1 | 1 | U |  | 1 | 0.32504 | 0.21952 | 0.34403 | 0.3454 | 0.25943 | 0.39595 | 0.94105 | 0.84615 | 0.86887 | 0.026386 | 0.072553 | 0.061048 | 56.8627 | 15 | -12 | 1.3726 | 0 | 1 | 1.6982 | 1 | global_median | 0
D | 2 | D2 | 1 | U |  | 1 | 0.28882 | 0.19046 | 0.31576 | 0.37309 | 0.29204 | 0.4142 | 0.77411 | 0.65218 | 0.76234 | 0.11119 | 0.18563 | 0.11785 | 53.7255 | 16 | -14 | -1.7647 | 1 | -1 | 2.2615 | 1.4142 | global_median | 0
D | 3 | D3 | 1 | U |  | 1 | 0.28882 | 0.18699 | 0.32504 | 0.40045 | 0.3258 | 0.43216 | 0.72123 | 0.57393 | 0.75213 | 0.14192 | 0.24114 | 0.12371 | 53.7255 | 17 | -15 | -1.7647 | 2 | -2 | 3.3338 | 2.8284 | global_median | 0
D | 4 | D4 | 1 | U |  | 1 | 0.31118 | 0.2011 | 0.34403 | 0.4273 | 0.36058 | 0.44976 | 0.72825 | 0.5577 | 0.7649 | 0.13772 | 0.2536 | 0.11639 | 55.2941 | 17 | -15 | -0.19608 | 2 | -2 | 2.8352 | 2.8284 | global_median | 0
D | 5 | D5 | 1 | U |  | 1 | 0.34403 | 0.231 | 0.37362 | 0.45349 | 0.39623 | 0.46697 | 0.75862 | 0.58299 | 0.80008 | 0.11997 | 0.23434 | 0.096867 | 58.0392 | 15 | -13 | 2.549 | 0 | 0 | 2.549 | 0 | global_median | 0
E | 1 | E1 | 1 | U |  | 1 | 0.33446 | 0.22714 | 0.35374 | 0.33543 | 0.25715 | 0.3824 | 0.99709 | 0.8833 | 0.92505 | 0.0012669 | 0.053894 | 0.033833 | 57.2549 | 15 | -12 | 1.7647 | 0 | 1 | 2.0283 | 1 | global_median | 0
E | 2 | E2 | 1 | U |  | 1 | 0.28012 | 0.18355 | 0.31118 | 0.36258 | 0.28869 | 0.4006 | 0.77259 | 0.63579 | 0.77679 | 0.11205 | 0.19668 | 0.1097 | 52.9412 | 16 | -14 | -2.549 | 1 | -1 | 2.915 | 1.4142 | global_median | 0
E | 3 | E3 | 1 | U |  | 1 | 0.28012 | 0.18014 | 0.31118 | 0.38939 | 0.32129 | 0.41852 | 0.7194 | 0.5607 | 0.74352 | 0.14303 | 0.25127 | 0.12871 | 52.549 | 17 | -14 | -2.9412 | 2 | -1 | 3.6947 | 2.2361 | global_median | 0
E | 4 | E4 | 1 | U |  | 1 | 0.29765 | 0.19397 | 0.32973 | 0.41569 | 0.35478 | 0.43612 | 0.71604 | 0.54674 | 0.75605 | 0.14506 | 0.26222 | 0.12145 | 54.1176 | 17 | -14 | -1.3725 | 2 | -1 | 2.6237 | 2.2361 | global_median | 0
E | 5 | E5 | 1 | U |  | 1 | 0.32038 | 0.21952 | 0.35374 | 0.44135 | 0.38904 | 0.45335 | 0.72591 | 0.56426 | 0.78029 | 0.13912 | 0.24852 | 0.10774 | 56.8627 | 16 | -14 | 1.3726 | 1 | -1 | 1.9708 | 1.4142 | global_median | 0
F | 1 | F1 | 1 | U |  | 1 | 0.33922 | 0.22714 | 0.35374 | 0.32467 | 0.25629 | 0.3676 | 1.0448 | 0.88624 | 0.9623 | -0.019043 | 0.052448 | 0.016691 | 57.6471 | 15 | -12 | 2.1569 | 0 | 1 | 2.3774 | 1 | global_median | 0
F | 2 | F2 | 1 | U |  | 1 | 0.27583 | 0.17677 | 0.30213 | 0.35122 | 0.28688 | 0.38569 | 0.78535 | 0.6162 | 0.78333 | 0.10493 | 0.21028 | 0.10605 | 52.549 | 16 | -13 | -2.9412 | 1 | 0 | 3.1065 | 1 | global_median | 0
F | 3 | F3 | 1 | U |  | 1 | 0.27583 | 0.17677 | 0.30663 | 0.37744 | 0.3184 | 0.40353 | 0.73079 | 0.55519 | 0.75987 | 0.13621 | 0.25556 | 0.11926 | 52.549 | 17 | -14 | -2.9412 | 2 | -1 | 3.6947 | 2.2361 | global_median | 0
F | 4 | F4 | 1 | U |  | 1 | 0.28882 | 0.18699 | 0.32038 | 0.40318 | 0.35072 | 0.42108 | 0.71635 | 0.53316 | 0.76086 | 0.14487 | 0.27315 | 0.11869 | 53.7255 | 16 | -14 | -1.7647 | 1 | -1 | 2.2615 | 1.4142 | global_median | 0
F | 5 | F5 | 1 | U |  | 1 | 0.30663 | 0.20836 | 0.32973 | 0.42826 | 0.38369 | 0.43827 | 0.716 | 0.54304 | 0.75235 | 0.14509 | 0.26517 | 0.12358 | 55.6863 | 15 | -12 | 0.19608 | 0 | 1 | 1.019 | 1 | global_median | 0
G | 1 | G1 | 1 | U |  | 1 | 0.33922 | 0.231 | 0.35374 | 0.31316 | 0.25686 | 0.35165 | 1.0832 | 0.89931 | 1.006 | -0.034726 | 0.046091 | -0.0025791 | 58.0392 | 15 | -11 | 2.549 | 0 | 2 | 3.24 | 2 | global_median | 0
G | 2 | G2 | 1 | U |  | 1 | 0.28012 | 0.18355 | 0.29765 | 0.33908 | 0.28658 | 0.36958 | 0.82614 | 0.64047 | 0.80539 | 0.082946 | 0.1935 | 0.093996 | 52.9412 | 15 | -13 | -2.549 | 0 | 0 | 2.549 | 0 | global_median | 0
G | 3 | G3 | 1 | U |  | 1 | 0.26317 | 0.17344 | 0.28445 | 0.36468 | 0.31713 | 0.38729 | 0.72166 | 0.5469 | 0.73447 | 0.14166 | 0.26209 | 0.13402 | 51.3726 | 15 | -13 | -4.1176 | 0 | 0 | 4.1176 | 0 | global_median | 0
G | 4 | G4 | 1 | U |  | 1 | 0.27583 | 0.18014 | 0.29765 | 0.3898 | 0.34837 | 0.40472 | 0.70763 | 0.51711 | 0.73545 | 0.15019 | 0.28642 | 0.13345 | 52.549 | 15 | -13 | -2.9412 | 0 | 0 | 2.9412 | 0 | global_median | 0
G | 5 | G5 | 1 | U |  | 1 | 0.29322 | 0.2011 | 0.32038 | 0.41428 | 0.38015 | 0.42183 | 0.70777 | 0.52899 | 0.7595 | 0.15011 | 0.27656 | 0.11947 | 54.5098 | 14 | -12 | -0.98039 | -1 | 1 | 1.7208 | 1.4142 | global_median | 0
H | 1 | H1 | 1 | U |  | 1 | 0.34886 | 0.2468 | 0.35374 | 0.30095 | 0.25886 | 0.33464 | 1.1592 | 0.95341 | 1.0571 | -0.06417 | 0.020722 | -0.024103 | 59.2157 | 13 | -9 | 3.7255 | -2 | 4 | 5.8206 | 4.4721 | global_median | 0
H | 2 | H2 | 1 | U |  | 1 | 0.29322 | 0.20471 | 0.30213 | 0.3262 | 0.2878 | 0.35236 | 0.89889 | 0.71129 | 0.85743 | 0.046293 | 0.14796 | 0.0668 | 54.5098 | 14 | -11 | -0.98039 | -1 | 2 | 2.4416 | 2.2361 | global_median | 0
H | 3 | H3 | 1 | U |  | 1 | 0.28445 | 0.19397 | 0.30213 | 0.35114 | 0.31746 | 0.36988 | 0.81008 | 0.61101 | 0.81681 | 0.091474 | 0.21395 | 0.087877 | 53.7255 | 14 | -12 | -1.7647 | -1 | 1 | 2.2615 | 1.4142 | global_median | 0
H | 4 | H4 | 1 | U |  | 1 | 0.28012 | 0.19046 | 0.29765 | 0.37562 | 0.3477 | 0.38716 | 0.74577 | 0.54777 | 0.76881 | 0.1274 | 0.2614 | 0.11418 | 53.3333 | 14 | -11 | -2.1569 | -1 | 2 | 3.1068 | 2.2361 | global_median | 0
H | 5 | H5 | 1 | U |  | 1 | 0.28882 | 0.20471 | 0.30213 | 0.39947 | 0.37839 | 0.40414 | 0.723 | 0.541 | 0.74758 | 0.14086 | 0.2668 | 0.12634 | 54.5098 | 12 | -9 | -0.98039 | -3 | 4 | 5.0952 | 5 | global_median | 0

## 05_REPLICATES_MEAN
ID | DF | Type | Conc | PAbs_Red_median | PAbs_Red_sd | PAbs_Green_median | PAbs_Green_sd | PAbs_Blue_median | PAbs_Blue_sd | L_median | L_sd | a_median | a_sd | b_median | b_sd | DeltaL_median | DeltaL_sd | Deltaa_median | Deltaa_sd | Deltab_median | Deltab_sd | DeltaE_ab_median | DeltaE_ab_sd | DeltaE_ab_chroma_median | DeltaE_ab_chroma_sd | CIELAB_ref_source | NReplicates | QCFlagged | QCCritical
1 | 1 | U |  | 0.11509 | 0.036042 | 0.21212 | 0.06126 | 0.10979 | 0.020109 | 55.4902 | 2.6164 | 15 | 1.4826 | -13 | 1.4826 | 0 | 2.6164 | 0 | 1.4826 | 0 | 1.4826 | 2.7295 | 0.89564 | 1.4142 | 1.2185 | global_median | 40 | 0 | 0

## 06_FITTING
Channel | FitType | n_points | m | q | R2 | RMSE | ID | DF | C0 | C0_sd | beta_k | bias_index_k
PAbs_Red | Calibration | 0
PAbs_Red | UnknownFromEpsilon | 40 |  |  |  |  | 1 | 1 | 1.0543 | 0.33017
PAbs_Green | Calibration | 0
PAbs_Green | UnknownFromEpsilon | 40 |  |  |  |  | 1 | 1 | 1.9432 | 0.56119
PAbs_Blue | Calibration | 0
PAbs_Blue | UnknownFromEpsilon | 40 |  |  |  |  | 1 | 1 | 1.0058 | 0.18421
L | Calibration | 0
L | UnknownOnly | 40 |  |  |  |  | 1 | 1
a | Calibration | 0
a | UnknownOnly | 40 |  |  |  |  | 1 | 1
b | Calibration | 0
b | UnknownOnly | 40 |  |  |  |  | 1 | 1
DeltaL | Calibration | 0
DeltaL | UnknownOnly | 40 |  |  |  |  | 1 | 1
Deltaa | Calibration | 0
Deltaa | UnknownOnly | 40 |  |  |  |  | 1 | 1
Deltab | Calibration | 0
Deltab | UnknownOnly | 40 |  |  |  |  | 1 | 1
DeltaE | Calibration | 0
DeltaE | UnknownOnly | 40 |  |  |  |  | 1 | 1
DeltaE_chroma | Calibration | 0
DeltaE_chroma | UnknownOnly | 40 |  |  |  |  | 1 | 1

## 07_METHOD_COMPARISON
Selected | Rank | Method | Family | ComparableGroup | CommonFactorsN | Score | ScoreFormula | RankMode | n_stdadd | n_unknown | Estimate_value | Estimate_sd | Estimate_source | expected_label_CRM | expected_id_CRM | expected_value_CRM | expected_sd_CRM | estimate_for_expected_CRM | delta_expected_CRM | recovery_pct_CRM | rel_error_CRM | expected_label_Colorimetry | expected_id_Colorimetry | expected_value_Colorimetry | expected_sd_Colorimetry | estimate_for_expected_Colorimetry | delta_expected_Colorimetry | recovery_pct_Colorimetry | rel_error_Colorimetry | BaseScore
1 | 1 | L | CIELAB | not_ranked | 0 | 0 | not_ranked |  | 0 | 0 |  |  |  | CRM | 1 | 1.92 | 0.01 |  |  |  |  | Colorimetry | 1 | 2.1 | 0.15 |  |  |  |  | 0
0 | 2 | a | CIELAB | not_ranked | 0 | 0 | not_ranked |  | 0 | 0 |  |  |  | CRM | 1 | 1.92 | 0.01 |  |  |  |  | Colorimetry | 1 | 2.1 | 0.15 |  |  |  |  | 0
0 | 3 | b | CIELAB | not_ranked | 0 | 0 | not_ranked |  | 0 | 0 |  |  |  | CRM | 1 | 1.92 | 0.01 |  |  |  |  | Colorimetry | 1 | 2.1 | 0.15 |  |  |  |  | 0
0 | 4 | DeltaE | DeltaCIELAB | not_ranked | 0 | 0 | not_ranked |  | 0 | 0 |  |  |  | CRM | 1 | 1.92 | 0.01 |  |  |  |  | Colorimetry | 1 | 2.1 | 0.15 |  |  |  |  | 0
0 | 5 | DeltaE_chroma | DeltaCIELAB | not_ranked | 0 | 0 | not_ranked |  | 0 | 0 |  |  |  | CRM | 1 | 1.92 | 0.01 |  |  |  |  | Colorimetry | 1 | 2.1 | 0.15 |  |  |  |  | 0
0 | 6 | DeltaL | DeltaCIELAB | not_ranked | 0 | 0 | not_ranked |  | 0 | 0 |  |  |  | CRM | 1 | 1.92 | 0.01 |  |  |  |  | Colorimetry | 1 | 2.1 | 0.15 |  |  |  |  | 0
0 | 7 | Deltaa | DeltaCIELAB | not_ranked | 0 | 0 | not_ranked |  | 0 | 0 |  |  |  | CRM | 1 | 1.92 | 0.01 |  |  |  |  | Colorimetry | 1 | 2.1 | 0.15 |  |  |  |  | 0
0 | 8 | Deltab | DeltaCIELAB | not_ranked | 0 | 0 | not_ranked |  | 0 | 0 |  |  |  | CRM | 1 | 1.92 | 0.01 |  |  |  |  | Colorimetry | 1 | 2.1 | 0.15 |  |  |  |  | 0
0 | 9 | PAbs_Blue | RGB | not_ranked | 0 | 0 | not_ranked | unavailable | 0 | 1 | 1.0058 | 0.18421 | unknown_from_calibration | CRM | 1 | 1.92 | 0.01 | 1.0058 | -0.91419 | 52.3859 | -0.47614 | Colorimetry | 1 | 2.1 | 0.15 | 1.0058 | -1.0942 | 47.8957 | -0.52104 | 0
0 | 10 | PAbs_Green | RGB | not_ranked | 0 | 0 | not_ranked | unavailable | 0 | 1 | 1.9432 | 0.56119 | unknown_from_calibration | CRM | 1 | 1.92 | 0.01 | 1.9432 | 0.023173 | 101.2069 | 0.012069 | Colorimetry | 1 | 2.1 | 0.15 | 1.9432 | -0.15683 | 92.5321 | -0.074679 | 0
0 | 11 | PAbs_Red | RGB | not_ranked | 0 | 0 | not_ranked | unavailable | 0 | 1 | 1.0543 | 0.33017 | unknown_from_calibration | CRM | 1 | 1.92 | 0.01 | 1.0543 | -0.86565 | 54.9138 | -0.45086 | Colorimetry | 1 | 2.1 | 0.15 | 1.0543 | -1.0457 | 50.2069 | -0.49793 | 0

## 08_LEGENDS
Term | Meaning | Formula | Unit | Where used | Shown when | Notes
a | CIELAB green–red axis | median ROI value after RGB → CIELAB conversion | dimensionless | RAW, SUMMARY | CIELAB enabled
b | CIELAB blue–yellow axis | median ROI value after RGB → CIELAB conversion | dimensionless | RAW, SUMMARY | CIELAB enabled
BaseScore/FinalScore | Intermediate/final score fields when present | same common score unless an explicit post-score adjustment is applied | dimensionless | METHOD_COMPARISON | method comparison | If identical to Score in the exported workbook, Score is the authoritative field.
beta_k | Slope ratio for one standard-addition curve | m_std / m_cal | dimensionless | FITTING, METHOD_COMPARISON | stdadd present | Used to assess whether StdAdd and calibration slopes are coherent.
beta_k/beta_mean | Standard-addition/calibration slope ratio | m_std / m_cal; beta_mean is the mean across available standard-addition fits | dimensionless | 06_FITTING, 07_METHOD_COMPARISON | calibration plus standard addition present
bias_index_k | Absolute relative slope bias for one standard-addition curve | ¦beta_k − 1¦ | dimensionless | FITTING, METHOD_COMPARISON | stdadd present | 0 means identical slopes.
bias_index_k/bias_index_mean | Relative slope-bias index | ¦m_std/m_cal - 1¦; bias_index_mean is the mean across available fits | dimensionless | 06_FITTING, 07_METHOD_COMPARISON | calibration plus standard addition present
BiasIndex | Relative slope bias between standard addition and calibration | ¦β − 1¦, where β = m_std / m_cal | dimensionless | METHOD_COMPARISON | stdadd present | Equals 0 when the standard-addition slope equals the calibration slope. Larger values indicate stronger slope bias.
C0 | Concentration in original sample | standard addition: DF × (q/m); unknown/CRM from calibration: DF × ((y − q)/m); unknown from epsilon: DF × PAbs/(epsilon × l_cm), converted from M to the selected unit | mM | FITTING, OVERVIEW | stdadd present or unknown/CRM present | Same symbol is used for standard addition and unknown/CRM results; FitType defines the method.
C0/C0_sd | Estimated original-sample concentration and associated uncertainty | standard addition: DF x q/m; calibration projection: DF x (y - q)/m; epsilon mode: DF x PAbs/(epsilon x l_cm), converted from M to the selected unit | mM | FIGURE_RGB.png, 03_OVERVIEW, 06_FITTING | standard addition, unknown or epsilon quantification present | FitType identifies which calculation path produced C0.
C0_sd | Uncertainty on C0 | Fit covariance for calibration/std-add; for epsilon mode, propagated PAbs replicate SD divided by epsilon × l_cm and converted to the selected unit | mM | FITTING, OVERVIEW | stdadd present or unknown/CRM present | Reported in the same concentration units as C0.
Channel | Fitted analytical descriptor | PAbs_Red, PAbs_Green, PAbs_Blue or diagnostic CIELAB/DeltaE descriptor | text | FITTING, METHOD_COMPARISON | always
CIELAB (L*, a*, b*) | CIE 1976 color coordinates derived from the RGB image | L* = 116 f(Y/Yn) - 16; a* = 500[f(X/Xn) - f(Y/Yn)]; b* = 200[f(Y/Yn) - f(Z/Zn)] | dimensionless | RAW, REPLICATES_MEAN, CIELAB diagnostics | CIELAB enabled | D65 is the reference white used by the sRGB/OpenCV conversion convention. References: CIE 1976 L*a*b* and CIE standard illuminant D65.
CIELAB_ref_source | Source of the CIELAB reference used to compute Delta variables | zero calibration or other reported fallback source | text | 04_RAW, 05_REPLICATES_MEAN | CIELAB/DeltaE outputs present | Reported only for CIELAB-derived variables because RGB pseudo-absorbance does not use this CIELAB reference.
ClipDelta | Positive calibration-derived clipping correction | max(0, y_expected − y_shifted) when deficit exceeds SD threshold | signal | FITTING | RGB calibration | Computed from calibration after S0 and zero-intercept fit; applied by same channel and same x.
ClipX | Calibration x values where clipping correction was evaluated | list of calibration concentrations | mM | FITTING | RGB calibration | Transferred to StdAdd at the same channel and same x values.
Code implementation | Main computational libraries and code provenance | custom Python code using NumPy, OpenCV, openpyxl and matplotlib | software provenance | all generated outputs | always | Sources/libraries: NumPy for least-squares solves; OpenCV for image masks/color conversion; openpyxl for xlsx export; matplotlib for PNG figures.
Col | Plate column index | 1-based column number | well-column index | RAW, REPLICATES_MEAN | always
CommonFactorsN | Number of common factors used in the score | 3 for R²_cal, R²_std, slope agreement; 1 for calibration-only or stdadd-only fallbacks | integer | METHOD_COMPARISON | always | Used to avoid comparing scores obtained from different formulas.
ComparableGroup | Set of methods scored with the same formula | calibration_plus_stdadd, calibration_only, stdadd_only, or not_ranked | text | METHOD_COMPARISON | always | Scores are directly comparable only within the same group. The group with the largest CommonFactorsN is used for Selected/Rank.
Conc | Configured concentration for calibration or added-standard wells | user/configurator input | mM | RAW, REPLICATES_MEAN, FITTING | always | Unknown wells may have undefined Conc.
Confidence class | Qualitative class derived from reliability score | HIGH >= 75; MEDIUM >= 45; LOW < 45; NOT QUANTIFIABLE when no valid quantification is available | class | OVERVIEW | always | Interpret together with the reliability reason.
delta_expected_*/delta_reference_* | Difference between estimate and external reference | estimate - reference | mM | METHOD_COMPARISON.png, 03_OVERVIEW, 07_METHOD_COMPARISON | reference values configured
delta_reference_* | Difference between estimated concentration and an external reference | C0 - reference_value | mM | OVERVIEW, METHOD_COMPARISON, FIGURE_RGB | reference values configured
Deltaa | Difference in CIELAB a* relative to the selected reference | a - a_ref | dimensionless | RAW, SUMMARY | CIELAB enabled | Reference source is reported in CIELAB_ref_source.
Deltab | Difference in CIELAB b* relative to the selected reference | b - b_ref | dimensionless | RAW, SUMMARY | CIELAB enabled | Reference source is reported in CIELAB_ref_source.
DeltaE_ab | Total CIELAB color difference | sqrt(DeltaL^2 + Deltaa^2 + Deltab^2) | dimensionless | FIGURE_CIELAB_DELTAE.png, RAW_DATA_DETAILS_CAPTION.txt, 04_RAW, 05_REPLICATES_MEAN, 07_METHOD_COMPARISON, diagnostics | DeltaE outputs present | Reference: CIE 1976 L*a*b* color-difference form.
DeltaE_ab_chroma | Chromatic CIELAB color difference without lightness | sqrt(Deltaa^2 + Deltab^2) | dimensionless | FIGURE_CIELAB_DELTAE.png, RAW_DATA_DETAILS_CAPTION.txt, 04_RAW, 05_REPLICATES_MEAN, diagnostics | DeltaE outputs present | Reference source is reported in CIELAB_ref_source.
DeltaL | Difference in CIELAB lightness relative to the selected reference | L - L_ref | dimensionless | RAW, SUMMARY | CIELAB enabled | Reference source is reported in CIELAB_ref_source.
DeltaL/Deltaa/Deltab | CIELAB coordinate differences from the selected CIELAB reference | DeltaL = L - L_ref; Deltaa = a - a_ref; Deltab = b - b_ref | dimensionless | FIGURE_CIELAB_DELTAE.png, 04_RAW, 05_REPLICATES_MEAN, diagnostics | Delta outputs present | The reference source is reported in CIELAB_ref_source.
DF | Dilution factor used to report original-sample concentration | C0 = DF x concentration in analyzed/diluted well | dimensionless | RAW, REPLICATES_MEAN, FITTING | always
Empty-well QC | QC based on empty-well robust SD and/or stored empty comparison | warning/watch/ok/not_available from drift score and robust SD thresholds | class | OVERVIEW | always | Flags comparability or background/empty-well drift.
epsilon | User-configured Beer-Lambert-like proportionality coefficient for PAbs quantification | PAbs = epsilon × l_cm × C_M | M-1 cm-1 | FITTING, OVERVIEW | unknown-only epsilon mode | C_M is mol/L. The analyzer converts the calculated concentration to the selected output unit.
Estimate_value/Estimate_sd | Representative concentration estimate and associated SD used for method comparison | derived from standard addition, calibration projection, or epsilon mode depending on Estimate_source | mM | METHOD_COMPARISON | method comparison
Estimate_value/Estimate_sd/Estimate_source | Representative concentration estimate, associated SD and source | selected according to Estimate_source | mM and text | METHOD_COMPARISON.png, 07_METHOD_COMPARISON | method comparison present
expected_*/reference_* | External reference metadata configured by the user | user/configurator input | text or mM | METHOD_COMPARISON.png, 03_OVERVIEW, 07_METHOD_COMPARISON | reference values configured | External reference values are checks only and are not used for ranking.
Family | Method family used in method comparison | RGB, CIELAB, DeltaCIELAB, or other | text | METHOD_COMPARISON | always
Field | Name of a metadata or overview item | label reported in the first column of the sheet | text | 02_METADATA, 03_OVERVIEW | always
FitType | Type of fit or concentration-estimation row | Calibration, StdAdd, UnknownFromCal, UnknownFromEpsilon, UnknownOnly | text | FITTING | always
flatfield_span | Image-level slow-field nonuniformity | (P95 − P5) / median of the slow grayscale field | dimensionless | METADATA | always | Part of rule-based image QC; larger values indicate stronger illumination/background gradient.
FracWellWarnings/FracWellCritical | Fractions of replicate wells with warning or critical optical QC | count / NReplicates | dimensionless | REPLICATES_MEAN | always
Geometry and epsilon/path-length quantification | Assumption used when epsilon-based unknown quantification is enabled | l_cm = (volume_uL / well_bottom_area_mm2) / 10; C_M = PAbs / (epsilon x l_cm) | cm and M | RESULTS_CAPTION.txt, RAW_DATA_DETAILS_CAPTION.txt, REPORT | epsilon/path-length mode configured | Assumes ANSI/SLAS-compatible flat-bottom microplate geometry; non-flat or non-certified geometries require separate validation.
ID | Sample or reference identifier assigned in the plate map | user/configurator input | text | RAW, REPLICATES_MEAN, FITTING, METHOD_COMPARISON | always
ImageWarning | Well-level optical QC warning flag | 1 if optical QC rules flag the well, else 0 | 0/1 | 04_RAW | always
ImageWarning_any | Group-level indicator that at least one replicate had an image warning | 1 if any replicate ImageWarning = 1 | 0/1 | REPLICATES_MEAN | always
IRLS | Iteratively reweighted least-squares robust linear regression with residual-based weights | minimize sum_i w_i (y_i - (m x_i + q))^2; w_i is updated iteratively from residual magnitude | dimensionless | FIGURE_RGB.png, FIGURE_CIELAB_DELTAE.png, RESULTS_CAPTION.txt, RAW_DATA_DETAILS_CAPTION.txt, 06_FITTING, 07_METHOD_COMPARISON, diagnostic fitting tables | always | Reference: Huber, P. J. (1964), Robust Estimation of a Location Parameter. Implementation: custom NumPy-based IRLS using repeated least-squares solves.
L | CIELAB lightness | median ROI value after RGB → CIELAB conversion | dimensionless | RAW, SUMMARY | CIELAB enabled
L/a/b | CIELAB lightness and opponent-color coordinates | median ROI value after RGB to CIELAB conversion | dimensionless | FIGURE_CIELAB_DELTAE.png, 04_RAW, 05_REPLICATES_MEAN, diagnostics | CIELAB outputs present
liquid_volume_ul | Liquid volume loaded into each well | user input | uL | OVERVIEW, FITTING | epsilon mode configured | Used only to calculate path length for flat-bottom wells.
LOD | Limit of detection | 3 × σ / ¦m¦ | mM | FITTING, OVERVIEW | calibration only or calibration present | Shown only for calibration rows.
LOD/LOQ | Detection and quantification limits from calibration | LOD = 3 x sigma_cal / ¦m¦; LOQ = 10 x sigma_cal / ¦m¦ | mM | FIGURE_RGB.png, FIGURE_CIELAB_DELTAE.png, 03_OVERVIEW, 06_FITTING, 07_METHOD_COMPARISON | calibration present
LOQ | Limit of quantification | 10 × σ / ¦m¦ | mM | FITTING, OVERVIEW | calibration only or calibration present | Shown only for calibration rows.
m | Slope of linear fit | y = m x + q | signal / mM | FITTING | always | Calibration and standard-addition fits use the same notation.
m/q | Slope and intercept of the linear fit | y = m x + q | response / concentration and response | FIGURE_RGB.png, FIGURE_CIELAB_DELTAE.png, 06_FITTING | fit rows present
MeanBG_* | Linearized local background intensity for the channel | background surface evaluated at the well and gamma linearized | dimensionless | RAW | always | Derived from inter-well background reconstruction.
MeanBG_Red/MeanBG_Green/MeanBG_Blue | Linearized local inter-well background intensity for each RGB channel | 2D background surface evaluated at the well and gamma-linearized | dimensionless | 04_RAW | always
MeanW_* | Linearized median well intensity for the channel | median ROI channel intensity after gamma linearization | dimensionless | RAW | always | Computed from selected well ROI pixels.
MeanW_Red/MeanW_Green/MeanW_Blue | Linearized median well intensity for each RGB channel | median ROI channel intensity after gamma linearization | dimensionless | 04_RAW | always
Method/Family/ComparableGroup/RankMode | Method-comparison identifiers and comparable-score grouping | text labels defining descriptor, family and score comparability | text | METHOD_COMPARISON.png, 07_METHOD_COMPARISON | method comparison present | Scores should be compared directly only within the same ComparableGroup.
n_points | Number of finite data points used by the fit | count of finite x,y pairs | count | FITTING | always | All finite points are retained by robust IRLS.
NClipPoints/ClipX/ClipDelta | Calibration clipping/baseline diagnostic fields | NClipPoints = number of adjusted points; ClipX = concentrations; ClipDelta = applied response adjustments | count, concentration, response | 06_FITTING | calibration present | Diagnostic only; clipping is not part of the common method score.
NReplicates | Number of replicate wells in a summarized group | count of wells in group | count | 05_REPLICATES_MEAN | always
NWellWarnings/NWellCritical | Counts of replicate wells with warning or critical optical QC | count over replicate wells | count | REPLICATES_MEAN | always
PAbs | Image-derived RGB pseudo-absorbance | PAbs = log10(I_BG / I_well) = -log10(I_well / I_BG) | dimensionless | FIGURE_RGB.png, RESULTS_CAPTION.txt, 04_RAW, 05_REPLICATES_MEAN, 06_FITTING, 07_METHOD_COMPARISON | always | Pseudo-absorbance is not assumed to be spectrophotometric absorbance.
PAbs_Blue | RGB pseudo-absorbance (blue channel) | log10(BG_blue/W_blue) = -log10(W_blue/BG_blue) | dimensionless | RAW, REPLICATES_MEAN, FITTING | always | Image-derived pseudo-absorbance with fixed full-background normalization.
PAbs_Green | RGB pseudo-absorbance (green channel) | log10(BG_green/W_green) = -log10(W_green/BG_green) | dimensionless | RAW, REPLICATES_MEAN, FITTING | always | Image-derived pseudo-absorbance with fixed full-background normalization.
PAbs_Red | RGB pseudo-absorbance (red channel) | log10(BG_red/W_red) = -log10(W_red/BG_red) | dimensionless | RAW, REPLICATES_MEAN, FITTING | always | Image-derived pseudo-absorbance with fixed full-background normalization.
PAbs_Red/PAbs_Green/PAbs_Blue | RGB pseudo-absorbance for the red, green and blue channels | log10(MeanBG_channel / MeanW_channel) | dimensionless | FIGURE_RGB.png, 04_RAW, 05_REPLICATES_MEAN, 06_FITTING, 07_METHOD_COMPARISON | always | Exported in standard RGB order even though OpenCV stores images internally as BGR.
path_length | Estimated optical path length | l_cm = (V_uL / A_bottom_mm2) / 10 | cm | OVERVIEW, FITTING | epsilon mode configured | Assumes 1 uL = 1 mm^3 and a flat-bottom well with nominal bottom area.
path_length_mm | Estimated liquid height | l_mm = V_uL / A_bottom_mm2 | mm | OVERVIEW, FITTING | epsilon mode configured | Converted to cm before Beer-Lambert-like concentration calculation.
Purpose | Short description of the sheet role | free-text description | text | 01_CONTENTS | always
q | Intercept of linear fit | y = m x + q | signal | FITTING | always | For standard addition, C0 is derived from m and q.
QCCritical | Group-level critical optical quality flag | 1 if FracWellCritical >= 0.50 | 0/1 | REPLICATES_MEAN | always | Critical flag; points remain reported and used.
QCFlagged | Group-level optical quality warning | 1 if any replicate has ImageWarning, any critical replicate, or median UsedFraction < 0.70 | 0/1 | REPLICATES_MEAN | always | Warning flag; points remain reported and used.
QCFlagged/QCCritical | Replicate-group QC warning and critical flags | rule-based aggregation of well-level warnings, critical wells and UsedFraction | 0/1 | 05_REPLICATES_MEAN | always
R2 | Coefficient of determination | 1 - SSE/SST | dimensionless | FIGURE_RGB.png, FIGURE_CIELAB_DELTAE.png, 06_FITTING, 07_METHOD_COMPARISON | fit rows present | Closer to 1 indicates better fit.
R2_cal/R2_std_mean/m_cal/m_std_mean | Calibration and standard-addition fit descriptors used in method comparison | R2_cal from calibration fit; R2_std_mean = mean standard-addition R2; m_cal and m_std_mean are fit slopes | mixed | METHOD_COMPARISON.png, 07_METHOD_COMPARISON | method comparison present
Rank | Workbook-level rank of each method | 1 = highest Score within the selected ComparableGroup | integer | METHOD_COMPARISON | always | Rows outside the selected ComparableGroup are not assigned a direct rank.
RankMode | Data basis available for ranking | calibration_plus_stdadd, calibration_only, stdadd_only, or unavailable | text | METHOD_COMPARISON | always
recovery_pct_* | Recovery relative to external reference | 100 x estimate / reference | % | FIGURE_RGB.png, METHOD_COMPARISON.png, 03_OVERVIEW, 07_METHOD_COMPARISON | reference values configured
rel_error_* | Relative error versus external reference | 100 x (estimate - reference) / reference | % | 03_OVERVIEW, 07_METHOD_COMPARISON | reference values configured
Reliability criteria | Rule used to flag a method as low reliability in method comparison | low if slope agreement < 0.5 OR R² < 0.8 OR ¦Δ¦ > threshold | rule | METHOD_COMPARISON | stdadd present | The Δ threshold is max(3×reference SD, 50% of the reference value), evaluated against the nearest available reference.
Reliability score | Overall report-level reliability score | starts at 50 and adds/subtracts rule-based penalties/bonuses for calibration availability, stored calibration, plate QC, critical wells, used fraction, empty-well QC, and ranking separation | 0-100 | OVERVIEW | always | Not a statistical probability; it is a heuristic audit score.
RGB linearization | Conversion of gamma-domain RGB values to an intensity-like scale | I_lin = (I/255)^gamma, gamma approximately 2.2 | dimensionless | PAbs calculation in 04_RAW and FIGURE_RGB.png | always | No automatic image correction is applied.
RGB to CIELAB | Diagnostic conversion from image RGB/BGR values to CIE L*a*b* descriptors | OpenCV BGR->Lab for 8-bit sRGB-like images; reference form: linearized sRGB -> XYZ using the standard sRGB-to-XYZ matrix and D65 reference white -> CIE L*a*b* | dimensionless | FIGURE_CIELAB_DELTAE.png, RAW_DATA_DETAILS_CAPTION.txt, 04_RAW, 05_REPLICATES_MEAN, DIAGNOSTICS | CIELAB/DeltaE outputs present | References: OpenCV color conversions; IEC 61966-2-1:1999 sRGB; CIE 1976 L*a*b*; CIE standard illuminant D65.
RMSE | Root mean squared error of the fit | sqrt(mean(residual^2)) | response unit | 06_FITTING | fit rows present | Interpret in the units of the fitted response.
Robust SD | Robust dispersion across replicates | 1.4826 × MAD, where MAD = median(¦x − median(x)¦) | same as variable | REPLICATES_MEAN | always | Computed within each replicate group.
Row | Plate row label | A-based alphabetical row label | well-row label | RAW, REPLICATES_MEAN | always
S0_calibration | Channel-wide low-signal offset estimated from raw calibration | max(0, −min(raw calibration response)) | signal | FITTING | RGB calibration | Estimated before forced-zero calibration and transferred to StdAdd/unknown data.
Score | Common method-ranking score | for calibration plus standard addition: SlopeAgreement^2 x sqrt(R2_cal x R2_std) | dimensionless | FIGURE_RGB.png, METHOD_COMPARISON.png, 07_METHOD_COMPARISON | method comparison present | Expected/reference values, recovery, SNR and clipping are not used in this score.
ScoreFormula | Formula used to compute Score | text descriptor | text | METHOD_COMPARISON | always | Documents which score formula was applied to the row.
Selected | Workbook-level selected method flag | 1 for the highest-ranked row within the most informative ComparableGroup, otherwise 0 | 0/1 | METHOD_COMPARISON | always | Selection in the workbook is derived from method-comparison rank and not from expected/recovery.
Selected/Rank | Selected method flag and rank position | Selected = 1 for the selected method; Rank = score order within the highest common-factor group | 0/1 and integer | 07_METHOD_COMPARISON | method comparison present
Sheet | Workbook sheet name | worksheet name | text | 01_CONTENTS | always
sigma_cal | Noise estimate for calibration | SD at zero concentration or median calibration SD | signal | FITTING | calibration only or calibration present | Source is reported in sigma_source.
sigma_cal/sigma_source/SNR | Calibration noise estimate, its source and slope-to-noise ratio | SNR = ¦m¦ / sigma_cal | mixed | 06_FITTING, 07_METHOD_COMPARISON | calibration present | SNR is diagnostic only and is not used in the common score.
sigma_source | Source of sigma_cal | median_calibration_sd, blank_zero_calibration_sd, or unavailable | text | FITTING | calibration present
SignalT_* | Transmittance-like intensity ratio before logarithm | MeanW_* / MeanBG_* | dimensionless | RAW | always | PAbs_* = -log10(SignalT_*).
SignalT_Red/SignalT_Green/SignalT_Blue | Transmittance-like ratio before logarithmic conversion | SignalT_channel = MeanW_channel / MeanBG_channel | dimensionless | 04_RAW | always | PAbs_channel = -log10(SignalT_channel).
SlopeAgreement | Agreement between calibration and standard-addition slope magnitudes | min(¦m_cal¦, ¦m_std¦) / max(¦m_cal¦, ¦m_std¦) | dimensionless | METHOD_COMPARISON.png, 07_METHOD_COMPARISON | calibration plus standard addition present | Equals 1 when slopes are identical in magnitude; lower values indicate poorer agreement between calibration and standard addition.
SNR | Slope-to-noise ratio | ¦m¦ / sigma_cal | dimensionless | FITTING | calibration only or calibration present | Used only for calibration rows.
Type | Well role in the analytical workflow | C = calibration; A = standard addition; U = unknown; other labels may be reported when present | text | RAW, REPLICATES_MEAN | always
UsedFraction | Fraction of ROI core pixels retained after filtering | n_used / n_core | dimensionless | RAW, SUMMARY | always | Lower values indicate stronger filtering. Closer to 1 means less aggressive trimming.
UsedFraction_median | Median UsedFraction across replicate wells in a group | median(UsedFraction) | dimensionless | REPLICATES_MEAN | always
Value | Value associated with a metadata or overview item | reported value for the corresponding Field | mixed | 02_METADATA, 03_OVERVIEW | always
Well | Human-readable well identifier | Well = Row + Col | well label | RAW, REPLICATES_MEAN | always
well_bottom_area_mm2 | Nominal flat-bottom well area used for path-length calculation | pi × (d_bottom/2)^2 | mm^2 | OVERVIEW | epsilon mode configured | Derived from the selected plate geometry.